Q. Would like to know what we can figure out about our crystal if it is not diffracting to high angles? Physical reasons and if it means that longer exposures can compensate for it?

A. It depends on the reason for the weak diffraction. To some extent increasing exposure times will let you improve the measurement, but the experiment may become unreasonably long. In cases where diffraction is too weak you might still be able to get a low resolution structure or failing that a unit cell. Trying a different crystal may be a better course of action.

Q. How can we start count k = 0, or how you decide or instrument decides?

A. We will cover this during our Data Integration lecture. Essentially you cannot chose this as there is only one correct solution.

Q. What was the difference between omega and phi rotation?

A. Omega and Phi are two different axes of the goniometer. Both can be utilized to collect diffraction data.

Q. What is the difference between omega and phi? Do we change them simultaneously when collecting data?

A. Omega scans are generally preferred for small molecule crystals because they cover a wider range of the Ewald sphere than phi scans. On the other hand, since the phi axis allows for faster rotation speeds, phi scans may be used when measurements can be completed with rotation around only one axis, such as in synchrotron radiation measurements of protein crystals.

Q. Can X-rays causes any changes in the material? if causes then how to avoid it?

A. X-rays are a type of ionizing radiation, which can cause radiation damage in materials/crystals.
One can try to change the wavelength and lower the temperature to limit the mitigation of radiation damage.
Furthermore one can utilize electron for diffraction, as they are very mild in terms of radiation damage.

Q. In a CAP window, you have I/sig and I/sigo. What is the difference between the two?

A. I/sig is the signal to background ratio. I/sig takes care of all measured reflections while I/sigo only takes care of the observed (intense) reflections.

Q. How does differences in Absorption Coefficient effect our experiment?

A. In crystals with high absorption coefficients, absorption prevents the inherent diffraction intensity from being obtained. This particularly degrades crystallographic statistical values and structural analysis results when the crystal exhibits significant shape anisotropy. To resolve this, software processing is performed to correct for the effects of absorption.

Q. What signal-to-noise range makes sense for a valid experiment?

A. The signal-to-noise ratio should be greater than 3 across the entire range up to a resolution of 0.83 Å. In CrysAlisPro, the measurement schedule is set to ensure the signal-to-noise ratio is at least 15 across all resolution ranges as default.

Q. What is the wavelength you use in micro diffraction?

A. Micro diffraction is typically conducted using Cu-radiation due to its superior intensity.

Q. What is the criteria for deciding the number of scan provide for the sample

A. The lower the symmetry, the more scans must be taken to achieve full completeness on datasets.

Q. After checking my crystal with NMR I note it has some impurity like starting material, is it possible I single out the needed crystal when I am measuring the crystal structure

A. Often materials crystallize as pure chemicals and do not form co-crystals or salts with starting materials of the synthesis. You can often separate your product easily under the microscope when selecting crystals.

Q. Is it possible to run for solving if we got to slow number of diffraction spots?

A. Small molecule crystal structures can typically be solved from <60% complete data. A resolution of ~1.2 Angstrom should be present.

Q. Is kappa-phi same with chi-phi goniometer?

A. kappa and chi goniometers have two different geometries. Kappa goniometers allow for higher completeness of data, important when measuring crystals in P1 space group.

Q. What is charge flipping method?

A. Charge flipping is a type of crystal structure solution method. Please refer to literature for a full explanation of the theory and algorithms used.

Q. How exactly does the detector's position affect signal quality? I'm referring to WIT, IVCR, etc.

A. The scattering intensity of the crystal is angle dependent. Low angle data is highly intense, high angle data is typically weak.

Q. What is the accepted value of I over sigma?

A. I/sigma value usually should be higher than 3 in all resolution ranges. CrysAlisPro set exposure time so that I/sigma >15 should be obtained as default.

Q. What is the role of braggs law in this concept?

A. Bragg’s law describes where we expect to see diffracted peaks.

Q. From the theta/2theta goniometry-scan, out main intension is to get crystalline peaks, What we get from phi and chi-scan?

A. We typically do not scan theta in single crystal experiments, this is normally done in powder diffractometry. Other scans are done to cover the data we need to measure for complete data .

Q. If we have 2d crystal growth material , is it the same for measuring?

A. Assuming you are talking about materials like graphene, it will depend on the thickness of the material. You will probably find diffraction is too weak if the material is very (atomically) thin and if the material is flexible this could also cause trouble for a normal diffraction experiment.

Q. How does the alpha particle change detection when finding the diffraction on certain metals when using Electron X-rays

Q. Can porosity affect the crystallinity of a sample? I have a material with a very good morphology, but it is highly porous, and I don’t obtain any diffraction peaks at all.

A. The ordered part of the crystal is, in porous materials, a fraction of the crystal volume. What’s more if the solvent is completely amorphous, it contributes to the background. This is a challenging situation.

Q. How to identify the crystal beam damage?

A. Diffracted intensity will irreversibly decline as the experiment progresses. You may also see visual changes to the sample, light dark patches or loss of clarity.

Q. What are its applications in thin film research, such as solid solutions used in high-efficiency solar cells?

A. Thin films are not typically handled with single crystal instruments .

Q. When working with sensitive compounds, are there specific crystallization technique or solvent systems you have found consistently more reliable?

A. We will cover crystallization in lecture 2.

Q. Is it a good practice to delete solvent molecule from the solved crystal structure If there is no any value added information with the solvent molecule?

A. It is only something that we do as a last resort. If the solvent can be identified, it should be included. At the very least it should always be reported.

Q. With a highly absorbent, weakly diffracting (tiny crystal) and suffering from beam damage such as lead iodide perovskites, are Molybdenum or Copper more optimal as sources?

A. Difficult but I would favour short wavelength with long exposures IF diffraction can be seen. This will minimise the absorption and likely the beam damage. If diffraction is too weak and the only way it can be seen is with copper, then that is the only choice.

Q. What does a grey contrast under crossed polarizers reveal about crystal quality?

A. Under cross polarizers you are looking for the light to be extinguished as you rotate the crystal. i.e. the crystal goes dim then bright then dim etc. If this is not happening it means you likely don't have a single crystal. It could be a twinned sample or it might not be crystalline at all e.g. glass. Generally twins will be bright and not go dim, amorphous materials might be dim all the time.

Q. What is mosicity actually!?

A1. https://en.wikipedia.org/wiki/Mosaicity.

A2. https://dictionary.iucr.org/Mosaic_crystal

You will find some good illustrations when searching for images of mosaicity -- blocks of 'perfect' tiny crystals that are arranged 'imperfectly' to form the actual macroscopic crystal.

Q. I have some crystals that are unstable to any oil we're using for mounting crystals. For now I'm placing the crystals on the microscope slide using a solvent and then mounting it on a loop without oil. Is there a better way to handle this type of crystals?

A. You could use a polar solvent or material in that case. Try using water, ethanol, glycerine or PEG in that case.

Q. I am using vacuum dry to remove the solvent inside the glove box. Is this effect the crystal formation?

A. While vacuum causes solvent evaporation and can precipitate crystals, applying vacuum is usually too fast and yields poorer quality crystals. You may experience different results depending on your chemistry.

Q. Is smaller the size, better the crystal in case of slow evaporation?

A. Generally smaller crystals exhibit better quality. Smaller crystals however may not diffract to high angles for a given diffractometer setup.

Q. When working with air or moisture-sensitive compounds, are there particular crystallization techniques or solvent systems that you’ve found to be more reliable?

A. In my case, I used layering technique in the schlenk tube with solvent mixtures MeOH-Et2O, CHCl3-MeOH, and also only MeOH for all the layers.

Q. In my case, my chemistry is sensitive to water and air. I used layering technique for my crystallisation process. Every time it's crystallise as crystalline powder or very tiny crystals. They are not enough to diffract well. Do anyone know how to improve the crystal size and quality for air and water sensitive compounds with any other methods?

A. I assume you are working using Schlenk techniques. Suggestion is to add a solvent in which your material is only slightly soluble and crystallize by lowering temperature in the freezer.

Q. My crystal is always look like fluffy I tried different solvent. Any recommendations?

A. It seems your material crystallized too fast. Try to slow down the process or enhance self-repair of the system, e.g. by crystallizing at higher temperature, slow cool down, addition of modulators (MOF), etc.

Q. How do you decide the right concentration of sample when starting a crystal growth experiment?

A. It is difficult to give an answer that is applicable for each case, as every system (crystallization method and molecule) is different.

Q. Is it advisable to use beaker for slow evaporation or small vial? Which one will be better?

A1. Both. Always put crystallization in a series.

A2. You typically want to reduce the surface are of the interface from solvent to air. Therefore the narrower the vial the slower the evaporation and the better the crystal, but this is not always true, or necessary. People often find NMR tubes are good for crystallisation for example.

Q. I would like to know about crystallization during heating curve, i.e. MOF and other hybrid molecules synthesis. Are there any good way to determine the parameter for better crystal growth?

A. Addition of modulators, sometimes in very large amounts, has proven to be effective in MOF crystallization.

Q. Hello, when using crystallisation by slow diffusion in gel, there is any recommendations about how can we improve the quality of our crystals (all my crystals contain twins)

A. Please have a look at recent publication by Bernhard Spingler et al.
https://doi.org/10.1107/S2052252525009716

Q. What exactly is the metastable zones?

A. A region of the phase diagram where the solution is stable and no crystallisation occurs unless the solution is in some way perturbed, e.g. nucleation sites manually introduced.

Q. How to crystallize hygroscopic or oily compounds?

A. Crystallizing hygroscopic or oily compounds is challenging because they resist forming ordered solid structures. Here are two specialized techniques you can try:
1. Co-Crystallization (using a "Crystallization Chaperone")
2. Crystal Sponge Method

Q2. How can we find a good seed for crystallization?

A2. A good seed crystal may be a crystal of your crude material. You can also try scratching the glass of your vial to create a surface your crystal to grasp on.
Certainly other materials could be tried as well, but should be crystalline. Some people even crystallize using a their own hair! Make sure to not use conditioner before.

Q. When obtaining crystals using the anti-solvent method, is it okay to add the anti-solvent until you see a precipitate beginning to form?

A. In that case you will generate seed crystals very rapidly and might end up with large crystals.
If you see precipitation form, add solvent until precipitate is dissolved and lower temperature of the solution slowly.

Q. I have some crystal but they are not in solvent, how can I use this to check crystal structure. Do I need to just add solvent into it or check the crystal structure like that?

A. You can measure single crystal diffraction data on your crystals.

Q1. Is there a H-tube which can use for the crystalization?

A1. I have used H-tubes for growth of MOF crystals.

Q2. Do you know where we can buy it?

A2. They may have been custom made by a glassblower

Q. Tried small molecule crystallization using evaporation method in room temp, it was having crystal in the solution stage but after complete evaporation it becoming flakes, why? how can we overcome this?

A1. Don't evaporate to completion, when you have nice crystals in solution stop the evaporation. It is almost always better to store crystals under mother liquor and fish them out for experiments, especially if they are delicate. Some crystals incorporate solvent in their lattice and losing it through over-evaporation destroys the lattice.

A2. May be due to collapse by solvent loss or rapid supersaturation at the end of evaporation.

Q. The crystal is unstable and the solvent may cause the crystal structure to collapse. Are there any good methods for detecting such crystals?

A1. Collect data at low temperature in crystallographic oil like paraton oil

A2. Try substituting solvent with higher boiling point equivalent solvent, e.g. dioxane instead of THF. 
Measurement at low temperature is needed to maintain solvent inside the crystals.

Q. regarding diffraction pattern is it always have to be concentric?

A. Diffraction patterns always have some symmetry to them, which arises from the arrangement of the molecules in the crystals.

Q. I tried p-chloranil in ethanol , slow evaporation, but instead of getting needle like crystals i got yellow, flakes?

A. One compound may have several crystal structure (polymorphism) depending on the crystallization conditions.

Q. Any tips to crystallize 'tricky' small compounds that behave like organometallic ones, such as riboflavin-Co2+ complex? They tend to form orange gels.

A. Sometimes it works to try 'in-situ' crystallisation -- try an excess of ligand with a tiny bit of metal (or the other way round). Set up many vials, with various compositions). And of course: water, methanol, ethanol etc.

Q. Can we use during evaporation method ( mixture of organic-inorganic) stirring under a temperature or that could affect the crystal growth when cooling occur , because i have noticed that during stirring , the crystal is very small and start to grow while there is adequate stirring speed but how can i take this crystal without being agglomerated later?

A. It is generally difficult to control the aggregation between crystals, but some additives like surfactants may prevent the aggregation. De-aggregation using sonication is also available when your crystals are mechanically stable.

Q. Is there any way to make highly dense pellets with high purity using solid state route method of highly moisturus and low melting point Compounds such as K2CO3?

A1. You may want to look at sample preparation methods for IR spectroscopy. Is melting an option?

A2. I want to make highly dense pellets for electrical measurements.

Q. I was working on dicarboxylic molecule (potential ligand for chiral cages/polymer). But it never crystallised in any solvent/s using any method, heating, pressure bomb, vapor diffusion, anti-solvent. Is there any platform, where we can check the feasibility/energetics of the anticipated cage/polymer crystals? 
Or anything other thing that can predict if any crystals can be formed at all (not Fancy computational studies)?

A. You could try to add a base (e.g. triethylamine) and form a salt. These often form crystals easily.

Q. Is the cubic shape of crystal better than the sphere shape for all the modern diffractometers? I heard, that some years ago the sphere shape was the best for the absorption correction.

A. A cubic (or spherical) crystal will always be beneficial for absorption correction. These days we have powerful routines to correct absorption effects even for plate or needle like crystals.

Q. Hello, what does mosaicity mean? Is there a limit number that should I avoid, and change for another crystal?

A. Imagine a mosaic tiling on the floor. The tiles may not be perfectly aligned slightly off. Now consider the tiles are small groups of unit cells in a crystal. Mosaicity is essentially describing how perfectly the tiles line up.

Q. I have crystal sample which grow nicely in freezer, but melt within 3 mins exposure to R.T. Is there any advice to mount that kind of crystal?

A1. In crystallographic oil at low temperature

A2. Move the fridge close to the diffractometer, load the sample in the fridge and quickly mount to the diffractometer at 100K, then measure the crystal.

A3. Yes, fast reflexes/

A4. You can cool crystals during selection under the microscope.
https://journals.iucr.org/paper?S1600576718003291
other models are available too.

Q. I had a crystal that were very pretty in appearance, they also diffracted very well, but the data was always shit. Why would that happen? I have attempted several crystallisation, and data collection, but it is same shit data despite prettiest crystal, and diffraction pattern?

A. Have you check if it is twin?

A2. Perhaps you are in the wrong symmetry. Double check it is not lower symmetry and twinned. It could also be merohedrally twinned. This makes the diffraction patter seem more symmetrical but unless handled in refinement, your model will be bad. Another possibility, unhandled disorder.

Q. For liquid-liquid diffusion, does the anti-solvent have to be miscible with the mother liquor?

A. Yes, the anti-solvent must be miscible with the mother liquor.

Q. Sometimes host is crystal but after adding guest crystallity is loose. Is there any technique to prevents from the loss of crystallinity ?

A. To prevent the loss of crystallinity upon the guest soaking, some factors are important including guest concentration, solvent dissolving the guest, evaporation speed, and temperature.

Q. What have you found to be the most effective way to improve analyte diffusion into a crystalline sponge when the guest is reluctant to go in?

A. To promote the guest inclusion, raising temperature and guest concentration are effective. However, some class of compounds can never be included due to their chemical properties or molecular sizes.

Q. Why larger detector area is not always better for better data collection?

A. The main advantage a large area brings is to measure more data at the same time which allows you to do experiments faster. If you are using the same sensor technology bigger is definitely better, but not all detector sensor technologies are created equal. If you have a large detector but it is noiser than a smaller one, the quality of your data may end up being worse or you may find weak data much harder or impossible to see. To compensate (which won't always help), you need to collect the data several times (redundancy) to improve the quality through statistical merging. In this case though the advantages of the larger area are lost because you have to do collect more data and experiments become longer. In fact often a smaller detector made which has no noise like our HyPix HPC detectors, it will perform better than larger detectors which are cheaper to make, but have drawbacks i.e. noise.

Q1. 1. Is it possible that the molecular configuration observed during the screening process is different from the final refined structure?
For example, could a structure initially appear as S during screening, but after full data collection and refinement, be assigned as R?
2. When using a Mo X-ray source for an organic compound, is it possible that differences in data collection or refinement could affect the assignment of R/S configuration, especially for molecules containing only light atoms?

A1. (1) The configuration of stereogenic centers in the molecules typically cannot change upon radiation. At beginning of measurement, only low data completeness is available and the Flack parameter cannot be determined reliably. Towards end of measurement, higher completeness of data is achieved and it may be that the correct handiness of molecule if different from the initial structure solution.
(2) Absolute structure determination of light atom structures (H C N O) should be done using Cu-radiation as the anomalous scattering and intensity is weak when using Mo-radiation.

A2. So, even if I use low-temperature (cryogenic) conditions with a Mo X-ray source, could this still affect the assignment of the R/S configuration?

A3. I have seen a poster in AsCA'25 that claiming even using Ag-source can determine the Absolute Configuration with Flack x -0.03(8), but using big crystal ~200micron, extreme data collection condition of Data Res 0.35Ang with 14-fold redundant.  As the anomalous scattering gets stronger at higher resolution.  Though Cu-Ka would be much easier.

A4. That’s a very good point. To be honest, I am still in the process of learning how to properly handle absolute structure determination, including how and where to refine the Flack parameter.
So far, my main focus has been on data collection strategies for natural product samples, especially optimizing crystal quality, temperature, and completeness.
I understand the importance of the Flack parameter for absolute configuration, but I am still developing hands-on experience in refining and validating it correctly. I would really appreciate any guidance or best practices on this.

A5. Nobody would employ Ag radiation to determine absolute structures of light atom structures. Cu is the way to go, or use electron diffraction. However, It may be a choice in case many heavy elements are present with sufficient anomalous scattering. Pure inorganics maybe!

A6. Sorry I should mentioned that the poster mentioned was using a crystal of organic compound with oxygen (-rich) atoms the heaviest atom, as a show case.  For my own experience, I have used Mo-Ka (while Cu-Ka is in service) to collect data with good enough Flack x for pure organics, but is quite challenge, not always succeed, big and high quality of crystal is need and carefully planning of data collection.  Cu-Ka is much straight forward.

Q. How or why Ag-target can performed high pressure crystallography ?

A1. Ag radiation has an even shorter wavelength than Mo radiation. When measuring High Pressure Cells, you want to have short wavelength to achieve more complete data and lower absorption. Please refer to our HP-Workshop on the user forum: https://www.rigakuxrayforum.com/forumdisplay.php?fid=120

Q. Which source may be optimal for really small crystals with big absorption?

A. Absorption coefficients up to ~40-50 mm-1 can be handled with Cu radiation by choosing a small crystal Rule of thumb is crystal size ~ 1/µ . Always choose crystal size according to wavelength/absorption coefficient and characteristics of the X-ray source.

Q. Why short-wavelength source (Mo and Ag) is more suitable for measuring charge density?

A. The shorter the wavelength, the more data can be measured according to Bragg equation "1/2 lamda".
In charge density analysis, you need to measure many reflections/data points to build up crystal structure model, e.g. in multi pole refinement.

Q. Cu k alpha has a higher wavelength as compared to Mo. In such cases if we have both metals and organic things crystallized together, that time won't that be difficult to resolve C-C bond which has a bond length of 1.54 A while using Cu K alpha ?

A. Cu radiation is perfectly suitable for measuring coordination compounds. You can measure data up to ca 0.79 Angström, C-C single bond is much longer.

Q. What is preferred to be used in organic-inorganic materials cu tube or Mo one or it depends on the size of molecules ?

A. Coordination compounds and MOFs can be measured best with Cu radiation.

Q. If we have some solvents trapped in the lattice and if we collect data both at higher and lower temperatures and we see the space group changing, does it mean that the crystal is going through a phase transition or it’s just the solvents that is creating trouble at different temperatures?

A1. A change of space group is a phase transition - just a different word. It may happen that the symmetry of the crystal is changed because molecules/solvent arranges differently inside the crystal

A2. So does it mean for a phase transition the whole structure has to change to call a phase transition? Just a change of the space group doesn't mean a phase transition?

A3. It is enough if the crystal structure changes partially.

Q. Why does the Flack parameter exhibit significant error when determining absolute configuration using Cu radiation? How can this error be reduced?

A. The higher the I/sig(I), the lower is Flack parameter. The higher the redundancy in the dataset, the lower is the error on the Flack parameter.

Q. Is the volume and cell parameters like a,b,c and angles changes/increases at high temperature measurement compare to low temperature.

A. Yes,  that is the norm: Unit cell parameters increase with increasing temperature. However there are also materials with negative thermal expansion coefficient.

Q. Is there a system with two radiation sources, for example, with two wavelengths, that we could use to remove fluorescence effects from some samples?

A. Yes absolutely! XtaLAB Synergy platform allows to integrate two different wavelength. Furthermore using hybrid photon counting technology (HyPix) you can suppress the fluorescence signal picked up by the detector.

Q. What type of glue can be used to fix the crystal to the loop? When should I use glue instead of just oil?

A. Normally I would use glue if you: a) want to permanently keep the crystal so it cannot slip or move or b) are performing room temperature or higher measurements in which cases the oil can be less viscous and not hold the crystal securely enough. As for the type of glue it depends a little on your use case, superglue (cyanoacrylate) or epoxy glue.

Q. In a case where your crystal moves from the loop, is it possible to observe this on your screen (after few frames) during the data collection?

A. We have some tools that can identify this in the data processing and data analysis side of things. There is an option on the lattice wizard "check for the sample jumping" which tries to find any sudden discontinuous changes in the diffraction pattern. The data processing has a similar option to recover the experiment by considering this type of change in the diffraction pattern

Q. I have microporous materials, which are usually weak diffracting samples, so I usually mount larger crystals to increase a bit the diffractions. Is this wrong? Should I focus on smaller xtals?

A. It all depends on absorption coefficient of the sample. You can certainly take a little larger crystal than the size of X-ray beam to increase scattering intensity in most cases.

Q. I must do my measurement in sealed capillary and it will be pressurized with gas (so not to thin). Are there procedures to account for absorption correction from cylindrical shape of capillary on top of "drown" crystal shape?

A1. The absorption correction we perform is called multi-scan. It uses symmetry equivalents to determine the changes in absorption as the crystal rotates in the beam. This approach is normally quite good at compensating for a wide range of cases like this where you have little choice. It cannot correct for data which is too weak to be observed of course. You can also (if you can see it) face index the crystal - i.e. describe its shape - which can help absorption correction algorithms by providing knowledge of how absorption physically should be. I assume this latter part is what you mean by "drown" crystal shape. Additionally CrysAlisPro interfaces with Ross Angel's ABSORB program which has some very powerful additional tools for absorption correction.

A2. Indeed- should be "drawn faces". CAP allows import of data  (frames) from beamlines/other vendors(sometimes better than native programs)- Is it possible to also add "crystal movie" and process that for absorption correction?

A3. That is more difficult, we need to know the microscope position and movie format etc. Because there is not 'standard' for this, it is not easy to support every combination and so we generally don't support those which are not Rigaku configurations.

Q. u >80 mm' recommended to switch to longer wavelength X-ray source. Why not shorter wavelength, I think absorption shall be reduced?

A1. Typically a shorter wavelength has a lower absorption. It depends on the absorption edges of the material.

Q. Yes, If my sample has a μ larger than 80 mm-1 under Mo , that means I need to try Cu？

A. if your sample has µ=80mm-1 with Mo, you need to take a small crystal. Maximum size is ~1/80 mm ~ 15 µm !

Q. Absorption coefficient with Cu should be much higher.

A. Yes This is what I mean, I think this sentence “u >80 mm' recommended to switch to shorter wavelength X-ray source.

Q. Once get a crystal of my sample but I never thought I will need to check the crystal structure so I filter off the solvent of crystallisation and keep the crystal in solid state. But I'm trying to recrystalised it again because I needed the structure unfortunately I can't get that done now. My question is how I can use the crystal that I have which is not in solution for crystal structure checking?

A1. If you have a crystal already, than measure a single crystal diffraction dataset out of it.

A2. I see. You have crystallised, dried the sample and now it is the only thing you have to work with. Well, you can still try to mount it onto the diffractometer with a loop and some mounting oil. If it doesn't diffract you are probably stuck. Your options are very limited, use a more powerful instrument than you have, use Cu radiation if you aren't already, try a synchrotron or you could crush it and use it for electron diffraction, if you have access to an electron diffractometer but they are quite a recent development and there aren't many around.

Q. Larger area CMOS/CPAD detector and smaller area HPC Detector work similar in terms of data collection? As the HPC will have less noise than CMOS detectors?

A. CMOS/CPAD and HPC detector have fundamental technological differences. You are correct that HPC is a noise free detector and CMOS will always have noise. HPC detector can be arranged in a curve around the sample to achieve a high theta coverage and shorten measurement times and improve data quality even further.

Q. which data is better to know the crystal structure by refinement, BT-XRD data or HR-XRD data?

A1. This appears to be a powder question. In this school we are only concerned with single-crystal diffraction. If I have misunderstood, please elaborate.

A2. Yes Sir, it is powder diffraction.

A3. Our school is for single crystal diffraction. If you have questions about powder, you should ask your local Rigaku powder team for assistance .

Q. Absorption coefficient is around 150 mm-1 for complex intermetallic. Those kind of crystal what you recommend and how to the issue

A1. This situation favours short wavelength Mo or even Ag, and a very small crystal. You also want the crystal to be as close to spherical as possible so e.g. cubes would be better than needles. Cutting the crystal can help make it more symmetrical or ideal, grinding into a sphere which is quite difficult and uncommon these days. You should definitely face index this sample for best results.

Q. We are selecting 20 micron size around, It is difficult to do face indexing and also we have only Mo facility only. Is there any other parameter correction will helpful. Do you recommend any specific strategy data collection?

A. For strategy, collect as much data as possible. More redundancy will help the multi-scan algorithm work as well as possible. You can try adjusting the spherical harmonics by hand in the refinaliser or equivalent stage for your vendor. You could also consider a smaller sample. It may still diffract well. If you can make its shape as symmetrical as possible, avoid rods, or needles, stick to cube type geometry or spherical. Try the Gaussian correction with beam profile enabled. It might also help.

Q. About sample centering if we use automatic gonio head, will the sample be centred automatically. the video  showed was about the manual gonio and centering the sample manually? if yes, why a manual gonio head is used anyways?

A. Yes, motorized goniometer heads can centre the sample automatically using optical image recognition or X-ray based centring. The reason manual goniometer heads exist is because they are more common, cheaper and still effective. Motorized heads are more expensive, a more recent development and they are bulkier which may affect strategy efficiency.

Q. Why do you sometimes see it is a 'well diffracting sample' but see a very few diffraction spots on the screen? Should you rely on it or always look for the sample showing a high number of diffraction spots?

A. This sometimes happens to me too. I usually try another crystal with more diffraction spots. If the diffraction spots are few and at low angle then it will be difficult to get a reliable structure and absolute configuration

Q. Is it safe to collect data overnight at room temperature? Would this pose any risk to the detector?
For context, the laboratory temperature is maintained at 22–25 °C with relative humidity around 36–45%.

A1. I have collected data for more than 1 week at room temperature and it was fine

A2. Concurrent single crystal diffractometers do not need constant attention and supervision. Data collections can indeed run multiple days, e.g. when running a charge density data collection.

Q. Are there any specific limits or recommended ranges for laboratory room conditions during data collection?
For context, our lab is maintained at 22–25 °C with relative humidity around 36–45%. Are these conditions considered safe for long or overnight measurements, particularly with respect to instrument stability and detector safety?

Q. Why does the reflection at the rotational axis continue to multiple frames of data collection?

A. This will be answered in tomorrow’s webinar. It has to do with the data integration and overlaps of reflections.

Q. Regarding closing the slits for twining crystals, Is there an ideal intensity value or should I play around until see the peak separation?

A. The advantage of a continuous slit is that you can close it just enough to achieve separation without sacrificing too much beam intensity. You should close only just enough to separate if your sample is weak. If it is strongly diffracting, close it enough to separate but don't worry too much about the intensity as long as data quality is sufficient

Q. Can we change the target resolution = 0.837 before start experiment?

A. Yes! You can change preferred parameters for the measurement on the fly

Q. How to calculate the absorption coefficient of a material? Could you tell me for an example like TiO2?

A. CrysAlisPro can calculate µ or you use external programs, e.g. https://11bm.xray.aps.anl.gov/absorb/

Q. I had instances where I change the strategies to have I/sigma and other parameters (individual I/sigma, merged I/sigma) and different thetas (because I read the value of 30 is usually good), but the data collection time changes without significantly changing the completeness. How does this work, what do the numbers there mean?

A. You may have missed to recalculate the strategy, or you measured with the wrong point group settings.

Q. I would suggest to collect data with resolution as high as possible (let say 0.81Ang for Cu) if the crystal can be diffracted that far, probably take a bit longer only.  As you can easily cut the data to lower the resolution but not easy to get higher resolution data while you need them.

Q. I would like to determine the absolute configuration of the compound; should the checkbox at this position （—Friedel mated are equivalent）be checked or unchecked？

A1. It should be unchecked. The information about absolute structure is contained in very small differences between the Friedel pairs. When we say all diffraction patterns are centrosymmetric, this is largely true except for these slight differences in the case of chiral materials.

Q. reason befind the radius 1/lamda, could you explain again... I think lamda coming from the source like copper or Mo.

A1. Yes, lambda is the wavelength: https://en.wikipedia.org/wiki/Wavelength

A2. https://dictionary.iucr.org/Ewald_sphere

Q. If we increase the exposure time, is it possible to increase the resolution? What might be the other way?

A. For a given crystal, crystal orientation and diffractometer, increasing the exposure time is the only way to increase the resolution.

Q. The dots we see in the screen are indicating single molecule or atoms in a single molecule?

A1. *All* atoms contribute to *all* dots. The *position* of the dots is determined *only* by the unit cell dimensions (a, b, c, alpha, beta, gamma). The intensity of the spots is determined by the contents of the unit cell (or rather, the asymmetric unit). Often, the asymmetric unit contains a single molecule.

Q. So, 1 dot is a reflection of all atoms and indicate whole molecule. Is it wright?

A. The dots are bragg peaks (we also call these reflections) The position of the bragg peaks tell you the shape of the unit cell and its dimensions. The intensity of the peaks tells you about the position of atoms within the unit cell. You need to do a bit of computation to get from the intensities to the atomic coordinates.

Q. What is the maximum permissible displacement (defined by distance standard deviation) for sample wobble? I wonder if the SD tolerance range also applies to macromolecules.

A1. First sample wobble is describing how badly the sample has been centred. I do not know the upper limits I'm afraid, but clearly if the numbers are large, you have performed your experiment poorly. You should seek to centre the crystal better and repeat (if it is preventing you from solving the structure.)

A2. Is there any rule of thumb for those categorized as 'large'?

Q. In many cases, I have seen the reflection at rotational axis continues to multiple frames. What is the reason behind this?

A1. Reflections close to the rotation axis are in the diffracting condition for longer. We apply corrections (lorentz polarization) to scale them. Very close reflections are typically unreliable and rejected because they corrections need to be very strong.

A2. I think,In single-crystal XRD, a reflection far from the rotation axis intersects the Ewald sphere only briefly and therefore appears in one or two frames, whereas a reflection on or very close to the rotation axis remains near the Ewald sphere over a wide rotation range and consequently appears in many consecutive frames.

Q. Is it goes like this?

A3. Yes.

Q. If the radius of the Ewald sphere increases by using a lower wavelength, then why do we say the Mo source gives lower resolution than a Cu x-ray tube source?

A1. You can collect to *higher* resolution with Mo than you can with Cu.

Yes, the intensity of Mo radiation is less than that of Cu (all else being equal) -- but there is a resolution limit for both radiations (due to the wavelength). You *can* get to higher resolution with Mo, but you may need longer exposure times to get there.

Q. Should we use the filters that are suggested (e.g. for P-lattice, F-lattice etc)?

A. Generally, yes. They will be properly activated depending on the crystal system you have from the indexing. The proper approach if wrong is to reindex the sample to find the correct unit cell. I cannot think of an occasion where I needed to manually change the filter or that it was helpful to do so.

Q. Is it safe and to what extent can one omit out bad reflections (what most crystallographers regard as junk) from the reciprocal space?

A1. By 'bad reflections' you probably mean reflections where their *measured* intensity does not match that of the *prediction* from your model. If your model is essentially correct (i.e. its predicted reflections generally match the measured values) and you then find that there are *few* bad reflections, then these are most probably 'bad' because of an error in the measurement -- they might have been under the beam stop, for example, or they might appear stronger than they should do because a cosmic ray happened to hit your screen just where the reflection was on the frame.

A2. Of course, a proper beam-stop mask will make sure that the regions in space that are behind the mask will not be taken into account.

A3. If you have *lots* of bad reflections, this could point to a systematic problem -- either experimentally (crystal or instrument not centered/aligned), it could point to a problem with the crystal (twinning, for example) -- and it could be due to highly anisotropic absorption etc

Q. Sometimes the diffraction spots in 3D Ewald explorer are generally well aligned, but I also observe some split reflections or weak spots around the main reflections, but it does not look like a typical twinning, since there is no clear second, well-aligned lattice visible. What could be the reason for these extra diffraction spots? Also, what would be the best way to handle or exclude these reflections during data processing or experiment?

A. First comment, if reflections are there then their cause should be understood before deciding to dismiss them. That said possible reasons from the top of my head:  1. you have mounted 2 crystals (passenger) 2. You have a cracked crystal 3. You have lambda/2 diffraction 4. You have superstructure (commensurate or incommensurate) reflections.

Q. When you click "Follow sudden changes of sample orientation" what values should you put in the orientation search and the search steps?

A. Normally the defaults are suitable. If they are not, you can try increasing the search range. I would normally leave the steps alone

Q. Normally the defaults are suitable. If they are not, you can try increasing the search range. I would normally leave the steps alone

A1. Like SADABS used for absorption correction, how I can use absorption correction in CrysAlis

A2. During data finalization you can apply absorption correction harshit

Q. what is the meaning of complete sphere and full sphere please explain it?

A. A full sphere is basically everything you can possible measure. Every accessible HKL index is collected. Complete sphere means you have collected them all.

Q. If CrysAlis chooses a non centrosymmetric but i think it’s not, can i change it? Or it’s best to solve it as non centro?

Q. Modelling disorder around multiple symmetry elements, for example, a mirror plane and a 2-fold axis within the C2/c space group. I appreciate if you can demonstrate it briefly.

Please show us how to add H atoms manually on olex2, i mean using commands.

A. Use HFIX in Mode menu on the top of the Olex2 window, select proper hydrogen type, and click atoms you want to add H atoms.

A2. To refine a model around symmetry element, which connects to itself as a result of symmetry generation, assign a negative PART number to avoid the self connection.

Q. 1. If a customer does not want to share the molecular structure model because the work has not yet been published, is it still possible to perform SC-XRD refinement?
How is this typically handled in a service crystallography context?
2. During refinement, is it possible for an initially assigned R configuration to change to S (or vice versa)?
Under what circumstances can this occur (e.g. data quality, anomalous scattering, and refinement strategy)?

A1. (1) The electron density map should reveal what the molecular structure is. There is technically no need to have a chemical drawing ready before doing crystallography, but it certainly is always helpful to have a rough idea what you are dealing with. Knowledge about the expected elements, solvents used during crystallization, precursors etc are all valuable information.
(2) It certainly happens frequently, that the initial assignment of R- and S- are not correct in the first steps of structure refinement. You can invert the structure model with command "inv -f"

Q. Do you mean this link?
A. https://rigaku.com/products/crystallography/techniques/absolute-structure-determination

Q. how to fix solvent disorder or any alkyl chain disorder

A. To fix disorders of highly flexible solvents or alkyl chains, use of geometrical restraints like DFIX and DANG will be helpful.

Q. Do we need to install ShelexTL or any other software before installing Olex2?

A. Olex2 can operate with Shelx programs but also comes with its own structure solution and refinement program.

Q. Is there any ShelXT refinement series in the forum?

A. Sheldrick's software can be downloaded from here: https://www.shelx.uni-goettingen.de/download.php

Q. How to remove this alert A: 
PLAT214_ALERT_2_A Atom C01J      (Anion/Solvent) ADP max/min Ratio        6.7 oblate
PLAT214_ALERT_2_A Atom C02P      (Anion/Solvent) ADP max/min Ratio        7.1 oblate
PLAT214_ALERT_2_A Atom C04J      (Anion/Solvent) ADP max/min Ratio        6.9 prolat

Q. for example i am aiming to make M-N-C type catalyst for its utilization against CO2 electrochemical reduction. I do use X ray absorption fluorescence technique (both in situ and ex-situ) to get to know about the structure type (symmetric/asymmetric), oxidation state of metal, coordination number, interatomic distances to know metal is attached with M/C/O. So the question is can we get the same type of information here?

A. Yes, single crystal X-ray analysis gives structure symmetry, coordination number, and interatomic distances as a shape of the molecule. From these information’s, you can also speculate the oxidation state of metals. However, SCX needs sample crystals, so in-situ measurement requires some techniques to crystallize the state.

Q. Sometimes due to symmetry I think, the carbons in a phenyl ring are not all visible, so I can't use afix 66 and I use dfix and dang to get it right, is this okay? Also, I have some cases where the phenyl rings are "cut" in the middle (and only a part of it is visible) and only when I use command grow (or control + G) the phenyl ring is formed. How can I fix the aromaticity of this type of rings when not all of it is visible in my asu

A1. Just fix the aromaticity on the atoms you can see and the ones you cannot see (ctrl+G) will be fixed automatically by Olex, as they are related by symmetry.

A2. It sounds like you are seeing phenyl rings on special positions. You can model these using 'PART -1' and 50% (or appropriate) occupancy.

Q. If you can see Q-peaks for hydrogens, is it better to change the Q-peak to hydrogen atoms or just add hydrogens geometrically?

A1. Both ways are acceptable, but the latter is usually used. This is because contribution of hydrogen to the diffraction intensity is very small, and the positions of the electron peak and the hydrogen nuclei are slightly different due to the withdrawal of the bonding electrons by the mother atoms (C, N, O, etc).

Q. What about the shift ? How can you improve it?

A1. By giving extra refinement cycles

A2. It’s not always the case if the structure doesn't converge.

A3. "Shift" can correspond to all parameters that you refine. E.g. atom positions, ADPs, extinction etc.
Eventually these parameters cannot be further refined, then the shift is zero.
If you have convergence problems, than you need to improve the model.

Q. If you get a really good fitting of Fobs vs Fcalc despite not having good I/sigma statistics can we still assert the crystallographic model as correct?

A. That depends on whether you meet IUCr guidelines or understand and can justify why you could not. Trying to refine against noise is not sensible and if your data is weak, there is more chance that you are forcing your expectation on the data, rather than letting the data speak for itself.

Q. How olex2 bypass squeeze works and should we have platon installed or linked with olex2 before doing solvent masking using bypass squeeze or both works separately?

A. Olex2 has its own algorithm. The technique originated in Platon. There are two options for achieving the same result. BUT I would prefer Olex2 as it is more streamlined and does not modify the original data file. The algorithm is trying to calculate the contribution of the 'missing'/'empty' parts of your model to the reflection file and subtract them out. It is possible to seriously bias your structure using this technique so be careful and describe everything you do and why. Use it sparingly.

Q. I had instances where the solution only works if solved in olex-2 solve, if any other method is used like ShelXT or ShelXS, it doesn’t solve it well. Why is this so?

A. We have different structure solution programs because one may work better for a particular material than the other. ShelXT works great for organics and coordination compounds. ShelXS works well if at least one heavy element is inside the crystal structure, e.g. inorganics. ShelXD for large organic molecules etc. Olex2.solve is another great one, can be used for many types of materials.

Q. What to do when there is no Q peaks around the atoms, however when the electron density Maps show there is something wrong around Cl and F atoms in a symmetric molecule? I have tried to fit all atoms possible, but still seems that something is wrong in the electron density maps. Could it be an occupational disorder? How to fix it?

A1. It may be a site disorder, where Cl and F atoms occupy the same positions as a disorder. Use EXYZ and EADP constraints to fix the Cl and F atoms at the same position, distribute their occupancies so that their sum should be 100%, and refine.

A2. Thanks! Should the atoms occupy 50% in each position? Or it is freely refined?

A3. Without special reasons, free refinement is recommended.

A4. If they are freely, the molecular formula seems weird. There is a way to edit it in Olex? In the cif file the formula should appear the as it is, I mean with point numbers in Cl/F, or can I consider whole number?

Q. In my structures I have strong q-peaks located very close to heavy atoms (Bi) and some in between bonds of heavy atoms. These values are around 8-5 in the Olex scale. I tend to chalk this up to heavy atoms being finicky but perhaps you can offer some insight into why this might happen. I am very certain that these q-peaks cannot be from foreign atoms. I am somewhat certain that they are not due to disorder since giving them a partial occupancy does not lead to an improved R1. For context this is a layered inorganic crystal with +/-2 charged layers, not a small molecule.

A. Did you consider careful absorption correction and choice of an appropriately sized crystal? It sounds your crystal was too large and you measured too small redundancy to employ empirical corrections effectively.

Q. In the Ctrl+M (Maps) mode, what do the green and red spheres mean? Does the red mean bad residual electronic density?

A. If you observe a large red residual electron density around an atom, it indicates that the atom you assigned is too heavy, and a lighter atom should be used to match the observed electron density. Conversely, more green density suggests that the assigned atom is too light, and an atom with a higher atomic number should be chosen to match the observed electron density values.

OlexSys has an excellent YouTube channel https://www.youtube.com/@olex2538

Q. Should we handle the disorder first and then add hydrogen atoms? Also, is it fine to deposit cif file that has some atoms without thermal anisotropies?

A. It depends case to case, but usually it is easier to assign hydrogen atoms after handling the disorder of the atoms they are attached to first,
Yes, sometimes you will not have enough parameters to describe all atoms with anisotropic displacement parameters. You will need to explain this in your checkCIF file before structure deposition.

Q. Why we lost the phases?

A1. When an electromagnetic wave is measured, only the square of the function is recorded. This erases the phase information. The only way to maintain phase information is to make use of a technique like holography.

A2. Yes, we can measure magnitude directly, not phase. We need to calculate what the phases were that gave rise to the magnitude. That is the whole goal of structure solution.

Q. How to solve and refine a structure where you have on an electron density two different type of atoms? For example there are two different metal ions on the same electron density

A. If two types of atoms occupies the same location in the crystal, please fix them at the position using EXYZ and EADP constraints.

Q. When both split disorder and positional disorder are present, unreasonable bond distances at that time how to fix the bonding distance

A. When the bond distances got unreasonable as a result of disorder treatment, use of geometrical restraints like DFIX and DANG will be helpful.

Q. Is it possible to fix a specific space group for a shelxt solution? For example, if I know that the unit cell should be centrosymmetric but the software gives me a chiral space group, are there command line options to force an achiral solution?

A. You can force ShelXT to use a specific space group with -s"SPACEGROUP"   e.g. -s"P2(1)_c" or -s"P-1" etc

Q. How can we know about this command?
like you mentioned here (Hadd - r).

A1. You want a list of olex commands. https://www.olexsys.org/olex2/docs/reference/commands-manual/all-commands/

Q. Sometimes the atom shows a negative value. I forgot is it called ndc? (or sometime like Negative real dispersion correction) Can you teach me how to resolve it?

A1. If you have an npd, it usually means you have the wrong atom on that site, or there is disorder or absorption is not properly corrected. There may be other causes too. You need to correct the model and/or ensure data collection and processing has been done well.

A2. What if I am sure that is the right atom. Then how can I resolve the disorder issue?

A3. Difficult to answer without seeing the structure. Each structure is a fairly unique puzzle to be solved.

Q. Can a PCR file from full proof program can be used in Olex

A. No, Olex2 cannot read PCR format.

Q. Sometimes we have more than one kind of solvents sharing common sites and disordered about the symmetry and even sometimes it’s not possible to know the exact nature of solvents present. Can we use mask in that case?

A1. You can assign multiple solvents/molecules/ions per void. 
This publication may come in handy as well: https://chemistry-europe.onlinelibrary.wiley.com/doi/10.1002/%28SICI%291521-3765%2819980615%294%3A6%3C1016%3A%3AAID-CHEM1016%3E3.0.CO%3B2-B

Q. I mean how to justify that what we have masked?

A. If a molecule cannot be modelled, than usage of solvent mask is fine. Best crystallographic practice is to use model the solvent, in some cases it is simply not possible.

Q. In which instances do you change the occupancy of certain atoms?

A. Right-click the atom, and use "chemical occupancy" to set a fixed occupancy.

Q. What will you suggest if only some solvent molecules can be welly located but others are in a mess.   Do squeeze for the rest of the mess or just erase all solvent molecules and squeeze?  How about ions?  We better not erase ions no matter how bad it is, right?

A1. Good practice is to model the molecules that you can really see and model. The others that only give diffuse electron density can be masked.
Certainly you can also mask ion, e.g. large soft ones like NEt4+ or there like.

Q. What to do if you are not sure regarding the solvent? Could it be as "unknown solvate"? Or would be better do not use mask at all?

A1. You should have other evidence. e.g. the synthetic chemist knows what solvents were used, recrystallization solvent should be known. Use the electron counts from the solvent mask to work out which (or which combination) might be present. You may need to make a judgement.

Q. What is hoft factor?

Q. Does olex2 software have any paralellisms wih Reitveld refinement?

A. Rietveld refinement fits a calculated diffraction pattern to powder data, refining an average crystal structure from overlapping reflections, whereas single-crystal refinement uses individual, resolved reflections from a single crystal to determine a much more detailed and precise atomic structure, including anisotropic displacement parameters and subtle disorder. Olex2 is a single crystal refinement suite.

Q. For the NoSpherA2 option, if the structure is not like a molecular crystal, how to use this option?

A. NoSpherA2 should not care. It is using non-spherical form factors that are calculated from your geometry.

Q. What is the difference between R1 and wR2?

A. https://www.shelx.uni-goettingen.de/shelxl_user_guide.pdf

Q. How to refine Hooft parameter if it is orange

Q. Can you explain the systematic absence plot a little? Sometimes we are getting blank window.

A. The plot shows reflections which violate the expected systematic absences and their intensity. These should be weak or it may indicate you are in the wrong space group or something else is going wrong. If you are getting a blank window, maybe you are in a space group which does not have any symmetry elements that cause systematic absences.

Q. What is SWAT?

A. Solvent water correction. It is a method of modelling diffuse solvent (Moews & Kretsinger, 1975)

Q. Is there a way to generate the crystallographic info as a table for the Supporting Information for journals? Thank you!

A1. Olex2 has the 'report' function -- where you can get a basic report that you can cut/paste. There's also the more sophisticated ReportPlus extension for Olex2, which makes complete reports

A2. Thank you very much, Horst!

Q. When merging data from "twins"/mutliple domains- is Rint still valid?

A. Absolutely. After merging a new Rint is calculated based on the merged data.

Q. How to open the file in crysalispro ?

A1. Generally datasets can be open by double clicking on either the *.par or *.run files. Please let me know which file you referred to a different file.

A2. When i open the .par or .run file it gives a diffraction pattern how to make structure from this. which command to tool i have to use for the same

Q. Many of my samples are medium-sized organic molecule (around 10kDa) and the crystals were often grown from water. And they often diffract to 1.0 Angs. Can they be published as you mentioned that the high resolution should be better than 0.84 Angs? The R factors are also high around 15% despite best attempts in refinement. Can they be accepted by IUCr standards?

A1. The 'rules' change for anything by 'small-molecule'. You are talking in kDa, so this tells me that these are not small-molecules!

A2. The molecules you describe should contain a vast amount of solvent/water. These type of crystals often have a resolution limit around 1.0 Angström, similar to porous materials (MOF, COF, etc) even when measured with strong rotating anode sources or at synchrotrons. 
Definitely such data is publishable, but with R1~15% it sounds you did not treat the disordered solvent correctly. Please refer to Horsts presentation on solvent masking.

Q. I would like to ask for clarification regarding CCDC data policy.
Is there any expiration or time limit for CIF files that have been deposited to CCDC but are kept on hold (not yet published)?
I am asking this to ensure I provide accurate information to users/customers.

A1. You best ask this to the folks at CCDC directly.
https://www.ccdc.cam.ac.uk/contact-us/

Q. Is it possible to quickly get CCDC numbers via OLEX2 ?

A1. There is a 'get ccdc number' button in Work > Report, but I am not sure whether this still works on the CCDC side. I think they want you go go through their portal instead.

A2. Got it. It seems that the users need to log in, deposit cifs and get numbers through their web platform

Q. When i open the .par or .run file it gives a diffraction pattern how to make structure from this. which command to tool i have to use for the same.

A1. you can go through them

A2. You need to ask your local crystallographer to do so. Or watch all the previous lectures :)

A3. This was demonstrated during lecture on Data Finalization.
You can conduct structure solution and refinement in the program of your choice (we recommend strongly Olex2 !). These programs have a direct interface to CrysAlisPro.

Q. Is measurement of density expected by CCDC or density only from refinement sufficient?

A. You do not need to fill in all available entries of your CIF file, but it is best practice to submit as much as possible.

Q. Can large synthetic macromolecules be deposited in CCDC as they cannot be deposited in the PDB?

A. Protein Data Bank (PDB) http://www.rcsb.org/#Category-deposit for polypeptides and polysaccharides having more than 24 units.

Nucleic Acids Data Bank http://ndbserver.rutgers.edu/ for oligonucleotides

https://support.ccdc.cam.ac.uk/support/solutions/articles/103000306409-where-do-i-deposit-data-not-accepted-by-the-ccdc-

Q. Usually I only address Alerts A and B. I usually ignore alerts C, is that okay?

A1. Not really -- alerts mean that that there is some sort of issue with the structure -- ideally, please look at all of them. If you can't solve the issue, please try and explain what the problem is and why it can't be solved.

Q. I can't install "Platon" on my OLex2? How can I install it?

A. Install Platon the normal way, and then let Olex2 know where it is by including the folder where it is installed in Home > Settings > Path

Q. How to install olex2

A. Please download an installer from the following URL:
https://www.olexsys.org/olex2/

Q. I think CrysAlisPro did a good job to prevent the loss of data at the low angle region due to overexposure by recollecting data with attenuation, while the losing of such data could cause Alert B sometimes.  
It that always true though?

A1. Data recollection was needed in the era of CCD and CMOS detectors. Concurrent HPC technology does not require re-measurement of overflowed reflections due to the superior count rate and counter depth.

So yes, if you have overflow reflections, then they should be re-measured with shorter exposure time.

Q. How could I justify this alert B: PLAT417_ALERT_2_B Short Inter D-H..H-D       H1       ..H3D      .       1.97 Ang.  x,y,z  =      1_555 Check. I have tried to fix the position of H, but at refinement cicle it moves a little bit.

A1. Due to misplaced hydrogen atoms

A2. You need to fix this. This means that H1 and H3d *in a neighbouring* molecule clash. In Olex2, right-click on H1 and then choose BANG (Bonds and AnGles) to 'grow' the arrangement.

A3. If you do BANG you can see the two hydrogens that are close and the use fit to move one of the hydrogens away

Q. When I get a Hirsfeld alert I usually use simu to get rid of the alert, is that okay?

A1. you should try and figure out *why* there is a problem.

A2. Usage of SIMU, DELU, RIGU can be justifiable case by case e.g. if the large ADPs are not due to split atom positions etc.

A3. Yes, that is assuming I have ruled out he case of disorder. Thank you!

Q. During distortion in phase change at higher temperature ,can the model be resolved to identify structure change at this temp

A. If the phase transition results in another crystalline phase it should be possible to solve and refine the crystal structure at the elevated temperature.

Q. Can a wrong space group cause to have that triangle in the structure?

A. Yes, wrong space group can cause the negative ADPs. Please try to check other possible space groups.

Q. Is this value of shift can be published in acta B ? (with desordred structure and twining ?)

A. Shift should be zero, maybe by fixing a certain parameter. If shoft still remains, than data or structure model must be revised.

Q. Could you briefly explain what is Hirshfelf Test Diff?

A1. https://journals.iucr.org/services/cif/checking/PLAT232.html

the components of the anisotropic displacement parameters along chemical   
 bonds are assumed to be equal in magnitude. Large differences might
 indicate contamination of these parameters with other (unresolved) effects 
 such as (substitutional) disorder, model or data errors and/or 
 over-refinement.

Q. May I ask how to show or what number shown ADP in Olex2 ?

A. Right click the background of Olex2, and select "Draw style" > "Ellipses & sticks".

Q. When depositing a CIF, should a disordered site be uploaded explicitly as a disorder model (e.g., two atoms with partial occupancies), or is it acceptable/preferred to represent it as a single atom?

A1. You only deposit a single CIF file. That file contains information of the disordered moieties as well.

A2. I understand you are trying to ask how to choose between modelling as a single atom with a large ADP or 2 disordered parts. It depends on how big that single atom is. I would also choose disordered parts when it described the chemical connectivity better and crucially refines stably. At some point you have to make a judgement on how big is too big. Also you should keep an eye on the data to parameter ratio. Obviously modelling disorder will increase the number of parameters you are modelling. You need to find the right balance.

Q. Sometimes we are getting alerts on solvent accessible voids even there is no significant electron density left to be model

A1. Yes, that happens.

A2. Is that means natural empty space in lattice gob ?

A3. Use solvent mask in that case. Interpretation must be done though.

A4. Ok thank you Gob

A5. You should write a comment for this alert. This could be due to desolvated solvate. You can find examples of comments for this sort of alert online from already published structures

Q. How does MicroED compare to an atomprobe?

A1. MicroED is a diffraction technique, whereas an atom probe is a spectrometry technique. During the analysis the crystallographer works out the arrangement of the atoms in the crystal by modelling.

A2. Thank you, so these are two different approaches for a similar study/outcome...?

Q. Can we determine the absoulte stereochemistry of a molecule without any heavy atoms (eg only C, H, N, O) using mED??

A1. Yes.

A2. Further reading: https://www.nature.com/articles/s41557-023-01186-1

Q. For powder samples in micro ED can you get the structure from grain by grain?

A1. micro ED is a single crystal technique. the sample is what we normally call a powder but we do experiments on individual grains which are just very, very small single crystals. the data we get is like a single crystal X-ray experiment not a powder experiment

Q. It seems the sample we need is actually what we use for powder XRD.  Then, will the size of the grains affect the result?  Still bigger the better or actually smaller is better (though typical size <100nm right?)? and more uniform the better or it doesn't matter.

A1. 3DED has an upper limit before absorption of the beam becomes intolerable. It is a balance. A crystal which is not too big, not too small. Thicker grains will suffer more dynamical effects (in general) so thinner is better from that perspective, but too thin may not diffract well enough.

Q. Is Rigaku XtaLab Synergy i capable of running this analysis or another instrument is needed?

A. No. XtaLAB Synergy-ED is our only electron diffraction instrument. The rrequirements are vry different from X-ray. XtaLAB Synergy-i, S, R, DW are all X-ray instruments.

Q. How is the vacuum-sensitive samples (porous MOFs with lattice solvents trapped inside) protected during ED measurements?

A1. Cryogenic sample holder for TEM is available for protection of such samples.

Q. Is it possible to compare MicroED and Precession Electron Diffraction?

A1. MicroED moves the crystal by rotation with a fixed beam, precession moves the beam by precession with a fixed crystal. You can get structures from both but the data you get is processed differently.

A2. Precession (PEDT) requires a higher dose and may not be ideal for radiation sensitive materials. MicroED works typically on doses < 1 e*A2

A3. If you do a tomographic series (discrete step-wise rotations) with precession ED (PEDT), you will get a very similar kind of dataset which you can likewise use for structure determination, though data reduction will need dedicated packages (like PETS2 or ADT3D). Nowdays, continous-rotation MicroED is used more often due to it speed and lower radiation exposure, but also PEDT has its advantages, e.g., it is great at cutting out tiny crystals within agglomerates or amorphous regions.

Q. Hello! Is the refinement similar to data from XRD? Regarding the Rint, generally is it higher?

A1. The procedure is virtually identical to single crystal Xray diffraction (and very different to PXRD). We use different scattering factors that better describe electron diffraction. In general R-factors are higher than we are used to from X-ray diffraction.

A3. Yes, all R-factors become higher because of the dynamical effects making each intensity different from the squared structure factors, and also to those that would be equivalent from symmetry alone. In refinement, this can be modelled, which will decrease R1, wR2, etc., ideally back to XRD level. On the other hand, during the data reduction, where you would investigate internal metrics like Rint (or rather Rrim etc.), this is not available, so the values are simply expected to be higher, which does no harm.

Q. The kinematic and dynamic methods, are they the different experimental methods or data treatment methods?  Which one is preferred and can they be done simultaneously and compare?

A1. Dynamic diffraction can also appear in X-ray diffraction but is typically small and can be ignored. In MicroED dynamic diffraction is the norm and can be treated during the data refinemet process. 
Olex2 or Jana2020 can allow for dynamic refinement.
It would certainly be best to treat each and every dataset with dynamic refinement but is usually only conducted when determining absolute structures. It furthermore helps finding hydrogen atom positions, in case they cannot be found using kinematic treatment.

A3. They are data treatment methods, in the final structure refinement step. Even in the dynamical method you will usually first do a kinematical refinement, so you will always have a comparison.

Q. When grinding the sample, how can you make sure you haven't lost crystalinity?

A. You don't normally lose crystallinity you are just crushing the large crystals into smaller crystalline grains but you could check with PXRD that you still have a pattern (if there is enough material) or just try it in the ED and see.

Q. Does ED encounter twinning problems?

A. It is absolutely possible, they are just very small single crystals after all. However since you have thousands of grains to choose from, you would normally not need to handle it. Just ignore the twinned grains for an easier life.

Q. If we have a crystal, good enough to perform electron density analysis upto 0.4 angstrom resolution.  By decreasing the  size of same crystal through grinding and performing electron diffraction. Which method should be preferably used, single crystal X-ray diffraction or electron diffraction?

A1. I assume you are talking about doing charge density with X-rays in your first sentence? As a general rule X-ray diffraction is the best choice if you have crystals that are suitable. It makes no sense to crush them for electron diffraction. Electron diffraction is for samples which do not work with X-rays. X-rays first, electrons as a fall back.

A3. Further reading on the topic:
https://journals.iucr.org/c/issues/2024/07/00/eq3017/index.html

A4. This sounds to me like a case where you'd just do single-crystal X-ray diffraction, unless you are looking for something that you need specifically electrons for (e.g., measuring the EDT instead of electron density - which may give you more information about charge states or protonation)

Q. What is the max tilt angle? Is "missing wedge" a problem for structure determination?

A1. The maximum tilt angle is typically limited by the geometry of the sample holder, e.g. +/- 80 deg.

A2. The missing wegde can be a problem, e.g. in low symmetry space groups like P1 and P-1. Here data of multiple crystallites can be merged.

Q. Does it matter which side of the TEM grid faces towards upcoming electron beam?

A. No, it really does not matter data quality wise.

Q. Collegue here is using Jeol TEM/STEM 200kV and tried collecting data like for structure solution in (rare) free time. Very complicated process of collecting and trying to process the data through train of programs for anything usable- ideally ins and hkl files for solution and refinement. XtaLAB Synergy-ED with CrysAlisPRo integrated package shown in the video makes it look like simple, accesible, straightworwad and fast thing to do.

A. Thanks, we are fortunate to have talented people working on these problems.

Q. Can microED/3DED also be  used to determine crystallite size?

A. Absolutely! The electron diffractometer still has imaging capabilitis.

Q. Will the refined result file still call .cif and I wonder will ppl still use checkcif to validate it?  or what to do in order to tell ppl the result is a 'reasonable' one?

A1. At the end of analysis you will have a .cif file at hand that checkCIF can process. CheckCIF recognizes that data is from MicroED and conducts structure validation accordingly. With MicroED come a few new entries that can be used, please refer to CIF dictionary. https://www.iucr.org/resources/cif/dictionaries
New entries are becoming available from time to time and CrysAlisPro it set up for this.

Q. Can CrysAlis-Pro be used to process data collected at the synchrotron?

A1. Yes! Please have a look at this webinar:
https://www.youtube.com/watch?v=BW_l5C8jAt4

Q. Could you please comment on how air-sensitive compounds should be handled when preparing microcrystals for ED experiments? In particular, it is not possible to use oil to protect the sample while transferring it to the diffractometer, right?

A1. There are special holders for the sample. There are cryo transfer holders which let you load the sample in a protective environment outside the diffractometer and freeze it before transfer to the diffractometer, but there are also air-free holders which can load the sample at room temperature without exposing to air. Oil is a bad idea, it will absorb the electron beam.

Q. Will the missing wedge angle problem limit the space group symmetry? Since low symmetry requires more data at different angles.

A. You can merge (combine) datasets of multiple crystals to overcome completeness issues in low symmetry space groups. You see it demonstrated just now.

Q. Just wonder where can I find the R(rim) in Olex2?

A1. In data finalization window, you can enable "Printout Options". It is located on the right side, rougly in the middle. I cannot share a screenshot here unfortunately. You can have a wide range of data quality indicators and also define more than 10 resolution slices for the data.

A2. Sure i will look for it.  
Though ppl are still focus on R(int) mostly, even during the checkcif validation process.  Comment have to be made for merged data (twin data)

Q. Is the data collection done for several crystals and merge the data, or for a good one crystal? In the video shows kind of screen for different crystals, is it just to screen and select a good one?  I could not get it.

A1. Different crystals. We are trying to stitch together a complete dataset and this requires crystals in different orientations so that we can access reflections which are not accessible in a given grain. This is a common techique in protein crystallography also but more because of beam damage destroying the crystals (which is also a valid reason in ED)

Q. PLEASE THE DIFERENCE BETWEEN THIS TECHNIQUE AND SYNCHROTRON

A. Synchrotron uses an X-ray beam just like a home lab instrument the only differences are that the beam is more intense, more parallel, is monochromatically purer, and can be tuned to a given wavelength. In this lecture we are talking about electron diffraction, this technique replaces the X-ray beam with an electron beam. Electrons also behave like waves and diffract from the crystal. Electrons interact more strongly with matter than X-rays so despite being lower power, strong diffraction can be observed.

Q. To obtain reliable absolute structure- are there some stricter limits on Rint, Rmerge, R and resolution?

A1. What matters most is redundancy and quality. You should ensure I/sigma is good and you have good redundancy. Higher resolution than you might normally collect might be helpful to improve data-to-parameter ratio. It is likely that a high R1/Rint etc. structure has an unreliable Flack, but as long as they are within normal ranges, you should be ok

A2. High I/sig(I) helps with the Flack parameter and high redundancy in the dataset helps with the standard deviation of the Flack parameter. So you need both, high intensity data and high redundency (or a heavy element with resonannt scattering). If you look at light atom structures (HCNO) the crystal must be chosen very well, as intensity differences of the Friedel pairs are low.

Q. What does e.s.d stand for in the previous slide?

A1. estimated standard deviation

A2. less than 2 esd in X means Fraser ?

A3. It means that the error on the Flack parameter should not be higher than twice the Flack parameter itself. Ideally you achive x = 0.01(2) or better

Q. If a compound has 1000 carbon atoms but only one heavy rare-earth atom, in this case, which wavelength, Mo or Cu is suggested for absolute structure determination?

A1. A crystal structure as you describe should not exhibit great diffraction strength and Cu radiation is the best choice in this case, as absorption will be low and resonant scattering is high.

A2. In this extreme case, is the only one heavy atom helpful for anomalous scattering  under Mo radiation?

A3. You should get a proper Flack parameter too.

Q. In monoclinic system space group P21/n why after solving get Z=2 because for monoclinic system Z=4.

A1. What is in the asymmetric unit? Do you see partial molecules? Sometimes you have a molecule lying across a symmetry element and you only see half of it. Maybe you have half molecules in your Asymmetric unit and we normally report the formula for the chemical moiety.  So when you expand the formula to describe the molecule, Z must be altered.

A2. Half molecule in the asymmetric unit.

Q. If I have Hooft Parameter -0.01(6) and a Flack Parameter that is essentilly meaningless (0.00(8) ), can I say that I can determine absolute stereochemistry?

A1. In that case you determined the absolute structure satisfyingly well. 0.00(8) is a perfect result!

A2. Remember the esd of 8 is on the last decimal place.

A3. Oh okay. Thank you guys!

A4. to make it even more clear:
0.00(8) means 0.00 +/- 0.008

Q. Is mean the value of F2+-F2-/mig higher the better?

A1. Ys

A2. Yes, the higher the easier it is to obtain a good Flack parameter. 
The value can be improved by measuring a high intensity dataset, adding high redundancy or a heavy atom like P, S, Cl, metal, etc

Q. are the Flack parameters of -0.04(8) or 0.03(8) acceptable to establish reliable absolute structures?

A1. For ab initio absolute structure determination Flack parameter should be smaller than x = 0.04(8).
If some chiral center is known, e.g. from synthesis, you can live with a higher value too.

A2. During CIF file preparation one would enter additional information:
E.g. Absolute structure was determined by anomalous dispersion, reference molecule, reference molecule + anomalous dispersion, synthesis, unknown, not applicable

Q. In the Ylid example now, the flack (-0.01(5)) for the second dataset is not good enough because the standard deviation (5) is > 2??

A1. i mean x > 2 as mentioned previously in the slides

A2. This is definitely good enough -- besides, 'Trust the Hooft Parameter'!

A3. Maybe a misunderstanding. x is the flack parameter, in your example -0.01, the esd is 0.05. The flack parameter should less then half of the esd value. detailed analysis by flack himself (from 2000)https://crystal.flack.ch/dox/ks0021.pdf

Q. This is a version of Olex2 specifically designed for ED? or it is normal one available from the olex2 website?

A. It's the normal version of Olex2, but the datasets need to be processed with a licenced version of CAP -- contact Rigaku for more details.

Q. How to reduce the hooft factor in refinement

A1. Basically, there's nothing you can do after the data collection. You may need to collect new data if you don't get a conclusive result

A2. So while strategy we have to increase the redundancy