Question 1 . What is the brilliance of the beam .

Answer . There is no one answer to this question. It depends on source type, wavelength and other factors like divergence slit settings.

Question 2 . How can one improve a poor R value of a crystal?

Does a poor R value have a serious implication on the credibility of the crystal stucture?

Answer . Yes a poor R value indicates the model is a poor fit to the data, you will likely struggle with publication if it is very high. There may be many reasons for this so it is hard to say how to improve it, a couple of tips, check the atoms are the correct type and none are missing , check the data was measured properly , don't takways trust the expected chemical formula, sometimes things can happen in the reaction you don't expect.

Question 3 . What is the actual accuracy of determining the unit cell parameters, the CrysAlisPro program gives results of 0.0001A, but when measuring the same crystal several times or taking different sizes of crystallites from one synthesis, the discrepancies are much greater at the level of 0.01-0.001A

Answer . It is generally known that modern diffractometers tend to produce very low esds on cell parameters. Differences between crystallites may be giving you some insight into the spread in quality of the crystals. You may also find this paper interesting… https://pubs.acs.org/doi/abs/10.1021/cg701260p

Question 4 . I have a crystal and its diffraction pattern is very good, but we can't get unit cells. What is the problem?

Answer . There could be many reasons for this. Most likely candidates are twinning , or incommensurate behaviour

Question 5. How the distance of the detector from the sample affect for the data collection?

Answer . Several things, longer data collections, more separated peaks, lower background if there is significant signal from things like flourescence, air scatter. There may be other factors.

Question 6. I work on metal-organic frameworks and sometimes I get better quality and size of crystals in 24 hours as compared to 72 hours. How can we explain that if crystal growth is directly dependent upon time?

Answer . Metal organic framework materials are unique.  These are froming more of a poylmeric crystal and so the standard things do not usually hold since you are trying to crystallize a material more all at once, not one molecule at a time.

Question 7. There are times when I am doing solvent layering, for example: dichloromethane and n-hexane, upon slowly adding the n-hexane, I get cloudy solution which will disappear after some time, is this common?

Answer. Yes, this is a good sign.  Stop layering, do not layer more hexane after it clears.   This is because the top layer of the DCM has more concentration of hexane and it is trying to ppt out.

Question 8.There are times when I am doing solvent layering, for example: dichloromethane and n-hexane, upon slowly adding the n-hexane, I get cloudy solution which will disappear after some time, is this common?

Answer. Thank you for the information and yes, I do get good crystals it I get the cloudy layer but I was thinking if this is common and why do I get the cloudy layer.

Question 9. Do these methods of crystal growth also works for needle crystals or only plate crystals?

Answer. Plate and needles differ in that it is which direction of the crystal grows faster.  Sometime change in solvent can effect crystal shape.

Question 10. Any time frame to allow the crystal growth?

Does it depend on the kind of solvent used, concentration, or any other factor?

Answer. Yes to all.  all of these factors affect the crystal growth of the material.

Question 11. Is growth of crystal depends on the outside environment? I mean outside temperature.

Answer. Definitely, where you pace the vial, is there vibration, warm, cold, light.

Question 12. I do get good crystals it I get the cloudy layer but I was thinking if this is common and why do I get the cloudy layer.

Answer. Cloudy layer is because, one your concentration might be too high or compound not very soluble in lower layer.  Place a layer of Toluene between them and this will prevent the mixing at the surface that causes the cloudy precipitation to form.

Question 13. Is there a duration required for crystal growth prior to SCX diffraction analysis or just depend on virtualization of the crystals in solution?

Answer. So the question is when are the crystals suitable for taking to put on the diffractometer.

1) Size, large enough to mount

2) Looks reasonable.

Question 14. Is your approach different if you grow crystals from a solution and from a suspension with already existing crystals?

Answer. If you are going to have crystals present, then most of the time you need the concentration such that not all the crystals will dissolve up.  There are two thoughts here, one is that the crystal provides a surface to grow larger crystals or two the presence of these crystals create the rest of the material to want to grow similar to these crystals.?

Question 15. Is your approach different if you grow crystals from a solution and from a suspension with already existing crystals?

Answer. The second one would be my case, powder crystal A + powder crystal  B + Water = The crystals I want.

But not both A and B dissolve in water.

Question 16. Apart from taking the crystals sooner, are there other ways to get smaller crystals (from the crystallization point of view) without crushing?

Answer. The best method I find for collecting crystals, is to use siliconize vials where the compound does not stick to the surface of the glass, and generally do not use evaporation.

Question 17. Of the typical crystallization methods, which would you recommend for the crystallization of mildly unstable cationic species? The quickest method?

Answer. Low temperature, saturat the solution and then place between -20 and-80 C.  Many crystals here at MSU have been done in pentane at -40 C.

Question 18. Any suggestion for "magic solvent" for compounds insoluble in any solvents? In my case it is often carbon disulfide (CS__2)

Answer. This is a difficult situation.  DMSO and DMF can possibly be used, but only small amounts of these solvents.  Good luck.

Nonpolar. Then DMSO/DMF does not help. Hex/Hept as well

Question 19. Do you have any tips for layering technique when using really polar solvents, such as water?

Answer. Water is a difficult solvent to work with to get crystals from.  (soluble hot insolugle cold is most used)  layering with alcohols may work, but usually very long alkychain alcohols are needed to get the compound out of the water.  Or  use Macro type method.

Question 20. We are preparing healthy fat from vegetable oils. Is it possible to form crystals in fat?

Answer. Your situation is very unique and also means that you need to purify as you crystallize.  I would say that low temperature is something you would want to attempt first.

Question 21. How can I siliconize a vial?

Answer. This is a standard method used in protein crystallization.  Check older papers to find a method.

Question 22. Is there a "universal" seed we can use to help crystallise a material?

Answer. no  I have had many people tell me of thing they added to the vial to get crystals, but nothing is standard at all.

Question 22. Determining the crystallinity of metal organic framework (MOFs) by SCX is often a big challenge often ends at PXRD analysis due to convoluted nature of MOFs, is there a way that your team were able to improve the determination of MOFs structure?

Answer. We crystallize these from the reaction mixture, by slow cooling of the solution.  Often a powder diffraction is all you can get, this is a tough compound to crystallize.

Question 23. Determining the crystallinity of metal organic framework (MOFs) by SCX is often a big challenge often ends at PXRD analysis due to convoluted nature of MOFs, is there a way that any team globally were able to improve the determination of MOFs structure?

Answer. Exactly. We tried several time to get the SC but still working with different modulators to enhance the crystallinity.

Question 24. Do you have any suggestions for improving crystallinity of a powder for microED, as opposed to getting single crystals for XRD?

Answer. This is a very interesting question.  As microED becomes more popular, this will be something people will be after.  I find doing the same type of things here, but of course since you are not worried about size you can afford to have it crystallize quickly.  Stir it in hot solvent?

Question 25. What are the methods that we can use to convert amorphous materials to crystalline materials in solid state chemistry?

Answer. Many people use anealing, Heating the compound up then letting it re-cool.  Some compounds need to get to melting point some do not.  This is a definite challenge.   Most I know about change the way the reaction occurs to try and generate non-amophous material.

Question 26. Is there any way to prevent twinning during crystallization?

Answer. Yes and no.  Often twining is an effect the crystal growth process and sometimes you can not change  the process, but a change of solvent often can help.

Question 27. How we can crystallize a macrocyclic host with organic guest, like pillarene calix arene?

Answer. Here is a paper in preferential crystallization of a calixarene and guest from a mixture. Maybe it helps? https://pubs.acs.org/doi/full/10.1021/acs.cgd.3c00007

Question 28. Some polymeric  compound  are not dissolved any solvent that time what i used?

Answer. it likely depends on the polymer length and type but not sure I can help here with this information.

Question 29. The solvent system is part of the reactant?

Answer. Not sure what this question means

Question 30. Is there any way to crystallize ionic compounds (ionic complexes) while ensuring there are no lattice solvents?

Answer. It may be difficult to 'ensure' this, try different solvent systems of varying polarity. You probably don't want to use something that will stick strongly, i.e. highly polar solvents. You may want to explore alternative approaches like microwaves and high pressure presses if you can't find a solvent system.

Question 31. Can you talk a little bit more about vial siliconisation?

Answer. Not sure what to add. Essentially, water contact angle to siliconized glass is higher giving a 'standing drop' which helps with vapour diffusion. It also prevents crystals from sticking to the glass.

Question 32. Which could be the best strategy to measure organic crystals with synchrotron radiation, with wavelength generally speaking between 0.4 - 0.1 amstrong?

Answer. Those wavelengths are fine because beams at synchrotron are very strong. So short wavelengths are fine anyway.

Question 33. you mean in terms of brilliance?

Answer. Yes

Question 34. Is there a way to measure the temp of the crystal directly instead of just the gas blown onto it?

Answer. I'm not aware of it. The cryo-systems on the market all measure the temperature of the gas coming out. Now ... the distance to the crystal is so short, the speed of the cold stream high and the crysal so small that it is assumed the temp of the crystal is really close to that of the gas blown out.

Not really, low temperature systems are very good.  I place a thermal couple in the diffractometer and center it and it reads a temperature within 0.5 degrees of the expected temperature, usually less.  I have done this with the oxford 800 and the 700.  The 600 the variation was almost 1 degree at times.

Question 35. Why should higher temperature get rid of incommensuration and twinning?

Answer. My best guess is that unit cell rearrange themselves as the temp increases. So they may yield to a toghter and more regular packing at higher temperature. Several crystallographers have told me that experienced such events ... sometimes. Worth a try.

I'm not an expert, but I was thinking about the electron density elipsoides which are elonagted at higher temperature, therefore, you'll have more overlap between the adjacent electron densities make them overlapped more leading to a single electron density (but not shperical) on your detector!

Question 36. How can the users better customize the order of the data collection runs? The CrysAlisPro program currently seems only allow users to sort out the runs based on the theta angles or keppa angles. But can users sort out the order of the data collection runs with more freedom in CrysAlisPro program? For example, in APEX program, the users can sort out the order of the data collection runs with any orders that they want to use.

Answer. The order of the runs is already optimize for (1) reaching completeness as soon as possible and (2) getting there as fast as possible. For instance, since omega can collect in both direction, a run might be situated to start where the previous run finished and collect this run counter-clockwise. As a result, I never change the order of the runs. But ... depends what you want to do. Why do you want to change the order of the runs exactly?

Question 37. For studying temperature-dependent phase transitions it's mandatory to use the same one crystal for HT phase and LT phase ? - or it's irrelevant and a number of different crystals can be measured.

Answer. I'd use the same crystal. The thing is that some transitions may take place if the crystal goes through a certain series of different temperature. I suppose you could use different crystals but make the new crystal go through the same temperature set points and for the same amounts of time. There is also the issue of histeresis. SO again, make sure all crystals go through the same series of temp set points and for the same amounts of time.

Question 38. Smaller crystals usually have longer measurement time. Is it not detrimental for the sample to expose it to Xray radiation for longer duration?

Answer. It may be, yes. But usually with small molecule crystals, radiation damage should not be your primary concern. If it does happen, then go ahead and use a larger crystal. Also, as I say in the presentation, I'm saying to keep experiment time reasonable, like from half hour to 2-3 hours max. DO not use a crystal that will require to increase experiment time from half an hour to 10 hours. Just use common sense.

Question 39. Smaller crystals usually have longer measurement time. Is it not detrimental for the sample to expose it to Xray radiation for longer duration?

I work with MOFs and in my observation, smaller crystals exhibited less redundancy with the default mode. However, on choosing complete coverage mode, the time varies by a factor of 4-10 times, changing it from 1 hour with 20-40% redundancy to 6-7 hour measurement with 90-100% redundancy.

Any suggestions on that?

Answer. Your low redundancy can be due to 2 reasons: (1) the symmetry of the lattice found upon indexing was higher than the true symmetry of your crystal. So the strategy was calculated with the wrong symmetry and the completeness was low at 40% or so. That would be why using 'complete coverage) improved your completeness. Or (2) your exposure time was too short and most of the high resolution reflections could not be collected.

Question 40. How to figure out my crystal is strongly absorbing?

Answer. The chemical content of the compound. Does it contain heavy elements? What are they and how many per molecule?

Question 41. How do I know if a compound/crystal has strong absorption?

Answer. Here are some options:

https://henke.lbl.gov/optical_constants/filter2.html

https://11bm.xray.aps.anl.gov/absorb/absorb.php

and the CAP formula editor.

For an absorbing compound the ideal thickness is that which gives a transmission of 50% along the longest crystal direction.

Question 42. Paratone N has been mentioned in the talk, it is quite a nice mounting liquid, though I think Hampton Research has discontinued the supply of paratone N for some time already.  Just wonder anywhere else is still available?

Answer. You are right, Hampton suggest this product as an alternative... https://hamptonresearch.com/product-Santovac-Cryo-Oil-633.html

I have no personal experience of it, but have used Fomblin oil before. I think any inert mineral oil of sufficiently high molecular weight should work.

A colleague also suggests that STP oil additive and mineral oil (Nujol) are also options.

I have tried the little bit pricy cryo oil but may not be as good as paratone N for many samples

Question 43. Where can I find the information about the size of x-ray beam for my instrument? I have a Rigaku XtalLab mini II.

Answer. The beam is 0.5 mm and that beam is flat, unlike the beam of the Synergy S used in today's presentation. So ... 0.5 mm or 500 microns.

Question 44. What are these horizontal and vertical shadows on blue screen?

Answer. The horizontal shadow is the little gap between the 2 modules of the detector. All detectors are a mosaic of module stuck together. It appears wide on the screen because the image is shown in a large scale. It is smaller on the detector.

The vertical shadow comes from the beamstop.

Question 45. For mosacity, there are three values: e1, e2, and e3. Should we look for smaller values of these three?

Answer. Yes, usually these values are between ~0.5 and 1 degree. Beyond 1 degree, it is quite a mosaic crystal.

Question 46. What is the shadow on the diffraction images (not the beam-stop)?

Answer. The horizontal shadow is the little gap between the 2 modules of the detector. All detectors are a mosaic of module stuck together. It appears wide on the screen because the image is shown in a large scale. It is smaller on the detector.

Question 47. How to use de-icing function? And when should we use it?

Answer. If your cryo-system works fine, you shouldn't have to use it. If it is not running so well, I'd use deicing for longer data collection of at least 3 to 5 hours. But that depends on how bad your cryo-system is. I think the feature is in the Options dialog in CAP.

Question 48. For some crystals, there seems no reflections at high theta area even very long exposure time is used. What should we do to such kind of samples?

Answer. (1) Try large crystal. (2) Go back to the bench and try to find crystallization conditions that will yield better crystal. (3) You can also try a synchrotron beamline where the beam's flux is really high.

Question 49. Could I know how to determine an optimal sample-to-detector distance?

Answer. It is done visually: you check whether neighboring reflections visually over lap or not i.e., they are a little on top of each other. But with Hybrid Photon Counting detector, you can collect almost ALL small molecule data at 30-35 mm.

Question 50. Could I know how to determine an optimal sample-to-detector distance?

Answer. Then use the closest distance possible ... unless you have a very long unit cell dimensions such as those occurring in protein crystals.

Question 51. Regarding choosing the proper radiation source, don’t you think that the presence of heavy metal nodes within the MOF lattice could affect the diffraction when Cu radiation is used? Just curious.

Answer. The ratio between the metal sites and all the other light atoms is usually low. Therefore, one can often use Cu for MOF crystals.

Question 52. I observed sometimes that many peaks fitted to my crystals but couldn’t find any space group. What is your opinion about this and what we should do with them?

Answer. Impossible to say without seeing the data. Could be a bad crystal, multiple diffraction patterns that confuse the indexing algorithm, or even a badly calibrated diffractometer where the detector distance or direct beam position is wrong.

Question 53. In which case is it better collet full data? And which case is better long redundancy data?

Answer. But you want a higher redundancy when (1) You need to get the absolute configuration of a compound, (2) you are dealing with a X-ray absorbing crystal. A higher redundancy will lead to a better empirical absorption correction which is done on symmetry-related reflections. And maybe (3) if you have a weak crystal that yield weak data at high-resolution, increasing the redundancy will improve a little the  I/Sigma at high resolution.

Question 54. Atlas's detector strategy says the image binning 2x2 4x4, what does that mean?

Answer. It puts together a grid of 2 pixels by 2 pixels i.e., making a larger pixel. This is the default binning for many compounds. You switch to 1x1 binning i.e., using each individual pixel as they are , when reflections are close together and you need to resolve them better. So typically, 1x1 binning is used for unit cell with large dimensions.

Question 55. Why large crystals, say 200um, are inappropriate for X-ray diffraction experiments, if the shape of our crystal is regular, e.g. a cube, you can introduce an absorption correction in CryAlisPro. What is the main substrate for which it is recommended to use crystals with dimensions of about 50um?

Answer. The smaller the crystal, the sharper the reflection.

Question 56. In the theory of static diffraction, the more planes involved in scattering, the narrower the diffraction peaks, so why is this so? Also the high-resolution reflections are weak, isn't this a disadvantage of using small crystals?

Answer. As crystals grow, unit cells will statistically adopt sub-optimum orientations, increasing mosaicity and thus decreasing reflections' sharpness. It is not always like that indeed, but usually. Also the longer the crystallization goes on, the more likely defects will statistically occur, again worsening the shape of reflections. As for experiment time, you are correct. Therefore, you must increase the exposure time. That's why I say in the presentation to use smaller crystals if diffraction from that sample is fairly good. You can't use small crystals if they already diffract weakly.

Question 57. The detector distance from the sample is related with the nature of the crystal or with the equipment setting?

Answer. Both. It depends (1) on the type of the detector and (2) on the unit cell dimensions of the crystal analyzed.

Question 58. Is there a reason you would ever change the detector distance?

Answer. Overlapping reflections coming from a long unit cell dimension (like a protein crystal) or a twin. If none of that, 30-35 mm for a Hybrid Photon Counting detector will cover the vast majority of your samples this is a different story with CCD and CMOS detectors that must be put a bit further back

Question 59. Is there any situation in which we have to use smaller or bigger scan width?

Answer. Not with modern diffractometers.

Question 60. What is acceptable I/sigma in terms of publication, 15?

Answer. 15 is good, But you may not always be able to get there if you have a weak crystal. Try to get at least 10.

Question 61. Would you suggest to always use the theta-dependency on the calculation of exposures on the collection strategy or just for low diffracting crystals?

Answer. Always use it.

Question 62. What is the difference of calculating for a 1 laue group or for a "sphere"?

Answer. Laue group 1 is really P-1. A sphere is P1.

Question 63. Why full sphere is more than double triclinic? Shouldn't it just be two times?

Answer. It should be about twice more. It depends how the strategy was calculated. Different 'seeds' can be used to calculate the strategy for a given crystal.

Question 64. What is your closet detector distances for all hypix type detectors (curve and flat ones)?

Answer. For flat HyPix, as close as possible to the beamstop. Usually 31-33 mm.

For curved detectors, strategies are better with a distance of about 40 mm.

Question 65. How to deal with very unstable crystals, which decay before completeness ~98% is achieved? is this possible to collect a few datasets (with completeness ~ 80-90%) from different crystals of the very same compound, the merge these datasets and to solve and refine the structure using this merged dataset?

Answer. The CrysAlisPro software allows to merge different datasets.

Question 66. During strategy collection it shows strong diffraction but very low I/Sigma. How will I increase the I/Sigma? The longer data collection could resolve the problem or something else. Please suggest.

Answer. Yes, longer exposure times.

Question 67. I have crystal which have very less size. it is a sheet type crystal but thin. Not coming in the loop. How to collect data?

Answer. You can also try using MiteGen micromesh loops.

Question 68. It seems that is not necessary the beamstop, is that correct?

Answer. That’s correct. As there is no air in the diffractometer which could cause background scattering, and the detector is radiation hard and does not have bleeding effects, there is no advantage in using a beamstop.

Question 69. How to load the sensitive samples?

Answer. If you mean samples which are vacuum sensitive, you would use cryo-transfer. Please visit this link https://www.youtube.com/watch?v=4Z32Sb5S1uA to see the full sample preparation for cryo-transfer and some detailed explanation. If your sample is air- or moisture-sensitive, you could prepare the grids in a glovebox, insert them in a small TEM grid storage box and proceed with cryo-transfer procedure like for vacuum sensitive samples.

Question 70. Given access to well-crystallized inorganic samples down to millimeters in size, what are the advantages of electron diffraction over X-ray diffraction? If it allows us to determine the absolute structure with better accuracy, how significant is this improvement?

Answer. Generally, for large crystals, electrons are not suitable due to their strong interaction. 3D ED is rather pertinent to samples which would otherwise not be amenable to single-crystal diffraction at all. Electrons can still yield some additional insight, such as charge states (due to scattering from the Coulomb potential) and a strong indication of absolute structure. However, for crystal sizes exceeding several microns, you’d rather want X-rays.

Question 71. What is the significant difference between TEM and this ED?

Answer. In terms of hardware, while the overall electron-optics and gun is similar to a TEM, the configuration is very much optimized for electron diffraction. Also, the type of detector is totally different from a TEM camera. Finally, the control system and user interaction with the hardware is very different and resembles more closely a SCXRD machine than a TEM.

Question 72. What is the purpose of doing GRID and what additional information do we get besides doing SCXD?

Answer. Different than in SCXRD, we do not just mount one single crystal due to the size, but thousands of crystals on a TEM grid. The grid is our sample holder. SCXRD has the 1-micron barrier. Samples smaller can't be measured with X-ray, hence ED is a nice new tool to get the crystal structure.

Question 73. What is the difference between the HRTEM and ED diffraction? Is it possible to collect the ED diffraction in HRTEM?

Answer. If you’re referring to the instrument here: ED data can be taken on high-end HRTEM machines, but it takes much more skill and effort, as those machines are not optimized for it.

Question 74. As an emerging technique, what are the main limitations of 3D ED technique compared to traditional SCXRD?

Answer. Certainly, some limitations are about the sample environment, e.g. the vacuum in the diffractometer - which can be worked around by cryo transfer. Another limitation is the unavailability of multi-axis goniometers, again due to spatial constraints and the requirement of vacuum transfer. This may necessitate merging of data from different grains, which is however straightforward.

Question 75. What is the beam size and the resolution of the ED?

Answer. On the Synergy-ED, the minimum beam size is 20 nanometers. However, we mostly use the selection-aperture approach, where the diffracting area goes as low as one micron, matching typical crystal sizes and the gonio’s sphere of confusion.

About resolution: if you refer to the diffraction resolution, this can be arbitrarily set by varying the detector distance.

Question 76. How is the integration process? Can the structure be solved in the traditional way without using Autochem?

Answer. Yes, absolutely. Virtually all standard crystallographic software will work with ED data.

Question 77. How long/difficult is cryo loading a sample in comparison to room temp?

Answer. Please have a look at the webinar: https://www.youtube.com/watch?v=4Z32Sb5S1uA . The sample preparation is more complex and it also takes longer for temperature stabilization. However, after doing the sample preparation a few times, it does not take much longer than with the ambient temperature specimen holder. In a normal working day, one can do 3-5 cryo samples. Standard RT measurement, maybe 6-8 samples.

Question 78. Do you have any example to be made by FIB machine without using grid?

Answer. Even for FIB-milled samples, you would first deposit them on a grid in a typical workflow. We do have examples for it. See e.g. in Truong et al. doi:10.3390/sym15081555

Question 79. ED resolution is determined by the distance like XRD. Unlike TEM determines the resolution by applied voltage?

Answer. I’d say that it is generally hard to compare the meaning of „resolution“ in ED and EM; in both cases, the achievable resolution is affected by a lot of factors.

Question 80. What is the advantage of ED compared to SCXRD?

Answer. In addition, bulk analysis. If you have a sample which has different phases, most likely you could catch it during screening and later measure whole datasets of the different polymorphs or phases.

Question 81. What are "acceptable" Rint values for ED? Especially in the context of merged datasets?

Answer. There are no „hard“ and generally accepted standards here. For datasets strongly affected by dynamical effects, values as high as 0.5 are still not indicative of bad data. That said, for typical organic samples, anything higher than 0.3 makes us suspicious.

Question 82. You just said that the R value usually decreased with increasing redundancy and that merging several grains would thus lead to a slightly lower R. Why is that the case? How are redundancy and R value related to each other?

Answer. R_int by definition will increase with higher redundancy, even if observations are perfectly consistent save for their noise. Hence, we recommend to rather look at R_rim, which adds a normalization factor that removes this bias (but not systematic increase due to dynamical effects)

Question 83. Is the wavelength of the ED changeable?

Answer. Yes, by varying the acceleration voltage.

Question 84. What does the gonometric and stereographic view mean/how are they helpful in Ewald Explorer?

Answer. Both are kind of legacy views and more common in minerology.

The gnomic projection can be used as alternative method for lattice indexation. Particularly for bad data, this is a great tool. You can look for "triangles" in the diffraction spots and index the cell with gnomic vectors.

Question 85. Are peaks that do not meet the Laue conditions for the determined crystal lattice also registered in the "peak list"?

Answer. Yes. All peaks that meet certain criteria are extracted from the frames.

Question 86. When we apply intensity and d value filters?

Answer. Intensity: for filtering weak features to find a lattice in the stronger reflections or filter strong reflections to find a weak superlattice.

D-value: for filtering out powder features like ice rings.

Question 87. Is there a way to remove ice or powder rings from your data besides just excluding all the frames where they're present?

Answer. use the skip d-value tool in CrysAlisPro:

right-ckick the frame, activate line-select, select the ice ring, righ-click and skip d-value

Question 88. I was wondering if modulated structure can be fully analyzed with CrysalisPro/Olex2. You showed that modulated lattices (q-vectors i presume) and their satellite peaks can be indexed and processed but I was wondering if is possible to fully describe modulation function(s).

Answer. Modulated structures can be processed in CrysAlisPro in combination wih Jana (2006 or 2020)

Question 89. After integration, the main screen shows only the results peaks.

Answer. Yes

Question 90. Can we use the Bruker apex data to integrate in riguku crysalis pro? and how?

Answer. You can. There is a frame import tool on the left hand side icons of the main window which guides you through the process. Note that the instrument model may need to be refined or preloaded from an known sample like the standard crystal (ylid) to achieve good results.

Question 91. Is there than any specific correction tips for needle like crystals or plates? To compensate for the shorter / longer direction?

Answer. I would recommend to cut the crystals to a length suitable for your setup. If this is not possible, you could try centering your needle in a way that you only measure it at the very tip.

Reduction of absorption can further be achieved by choosing a shorter wavelength, and higher redundancy in the dataset will help with emperical scaling.

Question 92. Won’t we lose information about any defects/anomalies or impurities if we choose the outlier rejection?

Answer. Typically you measure redundant data that means rejection of one reflection will not impact data quality due to symmetry reasons.

Question 93. Can the corrections that are used for Xray data directly be used also for electrons?

Answer. The empirical corrections employed for ED are practically the same as for X-ray data. Numerical absorption correction is unfortunately not possible.

Question 94. How we can filter the Rint?

Answer. One chooses the filter and a rejection criteria for it. Rint is just one of those and should be used with caution.

Question 95. In what cases is it better to use Analytical absorption (Clark & Reid) and in which Gussian grid (Numerical integration) correction?

Answer. Gaussian Grid allows for more grid points and beam profile correction, so that should be chosen by standard.

Question 96. If you have a crystal that has 2 possible space groups like C2/c and I2/a which one would you choose? And why?

Answer. The comment rule is to choose the space group setting with the smaller beta-angle.

Question 97. I am working on MOFs and the PXRD results shows sharp peaks which is an indication of crystallinity within the MOFs matrix. However, to get the SCXRD using XtaLAB Synergy in my school is what we are yet to achieve. The problem is that; It is not showing any diffraction after mounting the sample. Do you think using the concept of Rint filter and other correction parameters could improve the crystal quality?

Answer. Tim, if you have no diffraction on the images, there's nothing the software can do. A few things:

1- Is it possible the crystals decay when you take them out of their mother liquor to mount them on a loop?

2- Did you do something with the sample after you measured PXRD data? If the crystals don't show diffraction data, then they are no longer crystalline.

3- Is it possible that you have other species in your sample which are crystalline and they are the ones showing powder peaks? Remember that with PXRD, you measure data on a bulk sample, not just 1 crystal. What are the unit dimensions from the PXRD data? Are they consistent with a unit cell for a MOF sample?

4- Try longer exposure times or larger crystals.

Try Cu radiation if you are using Mo now.

Question 98. If sometimes the software changes from your initial unit cell to the niggli cell between the niggli and lattice screen, is there a way to stop this/a reason the software does this?

Answer. You can restrict CAP by choosing a "user cell" during unit cell indexation.

Question 99. How do I decide if it's a centrosymmetric or non-centrosymetric on a crystal?

Answer. The "E^2-1" value will indicate this. A further indication is the automatic structure solution from AutoChem and of course the data reduction tables and merging statistics. Crystal structure refinement will tell you again about being centrosymmetric or non-centrosymmetric.

Question 100. Which files in the data folder that record the data measurement detail, data integration and correction details?

Answer. The "Data Collection" tab in CrysAlisPro contains all necessary information, also you have access to the used run-list.

Also, for a lot of detail on the processing, the <sample_name>_red.sum file (it is a text file) can be studied. This file is also displayed by the "Inspect data reduction..." button.

Question 101. Does the face indexing always improve the Rint in the end?

Answer. Yes, the more you scale the dataset, the lower is Rint. You are not looking for the lowest Rint though when scaling your data!

Question 102. In Bruker software system, there is .ls file which records all details of the data intergration and .abs file which record all details of absorption correction. Are there similar files created by CrysAlisPro?

Answer. The "Data Collection" tab in CrysAlisPro contains all necessary information, also you have access to the used run-list.

Also, for a lot of detail on the processing, the <sample_name>_red.sum file (it is a text file) can be studied. This file is also displayed by the "Inspect data reduction..." button.

Question 103. How accurate does the chemical formula need to be for the absortion correction? What about if we don't know the exact composition?

Answer. If you choose the wrong absorption coefficient during scaling, you will generate an alert in CheckCIF. The easiest would be importing the formula from your refined crystal structure (.res file).

Question 104. How is it that we lose the phase information during data collection, leaving us with the ability to measure only intensity and not phase?

Answer. We can only measure the intensity of light as it hits a detector. We know that there *is* light, but we can't measure at which 'state of the wave' the light wave hits that detector. Light waves are just too fast!

Question 105. If in XRD or PXRD, we can determine our refinement is good through the chi square value between 1.0 to 1.5. How can we determine that our refinement results on this single crystal are good and valid? Are the parameters the same or do they have different parameters than XRD refinement in general?

Answer. I am not sure what the question is. Theare are many different indicators that tell us whether a structure is valid or not. Arguably, the most important one is the R1 factor, but wR2, GooF etc also tell us something. I think the most important indicator is the residual difference map -- Ctrl+M

Question 106. I have come across some cases where "add H" option dose not add hydrogen. Can be done in just situations?

Answer. You can always add hydrogen atoms geometrically if you know the relevant 'hadd' command: 'hadd 137' will add three atoms to a selected -Me carbon, or 'hadd 23' will add two hydrogen atoms to -CH2-CH2- sp3 carbonds, 'hadd 43' will add 1 hydrogen to phenyl rings etc. There are some YouTube videos on this on Rigaku channel.

Question 107. I have question about validation and goodness of refinement in single crystal with this powerfull software. If in XRD or PXRD, we can determine our refinement is good through the chi square value between 1.0 to 1.5. How can we determine that our refinement results on this single crystal are good and valid? Are the parameters the same or do they have different parameters than XRD refinement in general?

Answer. Theare are many different indicators that tell us whether a structure is valid or not. Arguably, the most important one is the R1 factor, but wR2, GooF etc also tell us something. I think the most important indicator is the residual difference map -- Ctrl+M

Question 108. How do you deal with it when it makes impossible bonds that shouldn't be there?

Answer. You can set the maximum number of bonds for any one particular atom (select it!) and then type 'CONN 4' -- this will then only show four bonds to that atom.

Or: "conn $Co 6" -- all Co atoms will have no more than six bonds. Or "conn $M 8" -- no metal atom will have more than 8 bonds etc.

Question 109. In general how many atoms (or may be the number of reflections) can be handled by olex2.solve (probably not as many as a small protein?)?  There is a limit for XS around 100,000 I can't remember the exact value and need to set the -b number somewhere for a very big structure, while getting a warning.

Answer. Hello Herman! olex2.solve is a very simply charge flipping implementation -- probably you'd use ShelXT! However, all dual space solution programs will get quite slow as the unit cell becomes big. ShelXS is probably still a good choice, or ShelXD, which (if I remember correctly) was made specifically for solving small proteins

Question 110. What information can we get from the R merge vs resolution plot?

Answer. How consistent equivalent reflections are at which resolution. At low resolution (measured near the center), they should be highly consistent (low R_merge). At high resolution, they will typically become less consistent -- but the better you can measure weak reflections (i.e. deal with the background properly, have enough diffraction etc) -- the better your R_merge will be.

Question 111. I was wondering if is possible to group lines in more than one group (drawn for highlight two differently oriented 1D-tubes in the crystal packing). When I try to do so and I generate the TIF file all the lines will be drawn with the same colour.

Answer. you can do any number of different lines styles by giving your lines names:

line -n=blue_line

you can then change the graphical properties of this line.

next time you want the same kind of line, simply use

line -n=blue_line

These lines should be remembered between Olex2 restarts, but can also be made part of a style.

Question 112. Why does the Rint not appear in the example?

Answer. Because the data have been merged after data processing -- something that shold really never ever happen.

Question 113. How do we know if there's solvent molecules in our crystal structure? And can I know how can I utilise the function of 'Use Mask'?

Answer. Well... if there is something that isn't connected to your main molecule, and is a solvent, then that's a solvent molecule in your structure. Sometimes it's not easy to figure out *what* that molecule is, though == solvent molecules can be highly disordered. They may also be counter-ions, of course.

Question 114. Why do the size of Q peaks automatically become smaller during the course of refinement?

Answer. They should not do that --but I've noticed this, too -- sometimes. It must have something to do with switching between 'pers' and 'anis' etc. I will need to investigate this behaviour. In short: it's a bug.

Question 115. Can we impose atomic ion scattering factors for olex2? e.g., when solving the structure of NaCl, we want the program to take into account f_calc atomic scattering factors for Na+ and Cl-, respectively, is this possible?

Answer. Yes, these are tabluated and available -- but I'd have to look up the syntax of how to include them.

Possibly a good way to really get to the bottom for these sorts of things, you could consider using non-sphercial form factors (e.g. NoSpherA2) for the refinement.

Question 116. Why there is red region around cadmium in this particular structure?

Answer. The absorption correction hasn't been done properly in this dataset.

Question 117. In the case of highly disordered atoms, sometimes when I try to add the Hydrogen atoms, olex gives an error about the number of bonds. For example, with tetra butyl ammonium, one of the terminal atoms is very disordered. It is an sp3 with 3 H atoms, but a cannot add the H atoms in both disordered carbons

Answer. The most common reason for this is that the disordered bits haven't been put into the correct PART. Often (and I suspect that's what happened in your case), the whole thing needs to be in PART -1.

Question 118. Hi Prof Puschmann, why Olex2 uses 3sigma rather than 2sigma of I/Isigma(I) to determine if the signal detected is basically noise? Can we change that in Olex2 to let the program to use 2sigma?

Answer. If you mean the line in the display of the plots: that is just a line; it doesn't actually do anything. 2 sigma *is* used for calculations, for example:  R1 was 0.0258 (I≥2 sigma(I))

Question 119. What is a good R1?

Answer. That very much depends on the system, but the colours in Olex2 are a good guide for 'standard structures'. Green is good, orange not so much and red really needs to be justified. For example in a very large cell, R factors can be higher than one would expect.

Question 120. How to correct "warning : 0 atom may be split and 1 atoms NPD" error during refinement ?

Answer. You should never really get N.P.D atoms. If you do, you may have a wrong atom type there. Twinning could also be a reason for N.P.D atoms.

Question 121. How do you explain the negative potential around the Cd?

Answer. The absorption correction was probably not done well in this dataset.

Question 122. Where do I find all the comands for Olex2?

Answer. Please visit olexsys.org, especially this link: https://www.olexsys.org/olex2/docs/reference/commands-manual/all-commands/

Question 123.XT is very powerful though I found sometimes it failed to solve high symmetry structure while XS gives the answer straightly (so just try all program available, in the old day I just use BLOC to deal with the very big structure and matrix -b can be adjust to include more reflections.  Dose olex2.refine has such limitation?

Answer. No -- there is no such limitation in olex2.refine. However, in the current version, it can become too slow to be practicable. In the next update this will be fit, and olex2.refine will also be suitable for large structures (already out in the dev version)

Question 124. Is olex 2 able to handle Electron Diffraction data?

Answer. Yes. It works well.

Question 125. How can the water molecule be refined? The ellipsoid is large.

Answer. It is free to 'move about' -- sometimes, refining it as disordered may help --otherwise restraints could be used. However: I'd just leave it as it is: the large ADP carries some information, after all!

Question 126. What can a high value of the GooF parameter (above 1.2) mean with relatively low values of R1 and wR2 (below 4%), should we worry about it?

Answer. This usually means that the weighting scheme has not been switched on.

If it *has* been switched on, and GooF is still showing as large, then there is something strange going on with the e.s.d. values of your measured reflections.

Question 127. Is there a way to visualize the symmetry elements in olex2 ?

Answer. No. We are not currently planning to implement this. However, if you want to know if an atom is on a special position, simply select it and type 'degen':

Mn1 [+X,+Y,+Z] 2

-X,1-Y,-Z

Site constraints: X=0, Y=0,Z=0

Or press CTRL+G, and you will see 'growing bonds' if it grows.

Question 128. Does OLEX2 automatically give the site occupancies of the disordered atoms?

Answer. Yes, but you can set them manually too: 'part -1 10.5'

Question 129. That, R, Rwp and GooF give us that information. Based on your experience is there a certain value parameter for some of these indicators to be considered good?

Some journal literature that I read for R or Rwp is said to be good if the value is below 10% but this is for refinement on XRD Pattern, is it the same in this case?

On the other hand, for this GooF value, I have not found any literature related to what number of GooF indicators is said to be good.

Answer. GooF should be 1. Depending on the size of the e.s.d of a reflection, it is either up- or down-weighted. I.e. if we know that a reflection has been measured badly, then we don't give it as much weight in the refinement as we give a reflection where we know it's been measured well.

Question 130. Could you please explain the importance of selecting or deselecting EXTI?

Answer. If the physical phenomenon of extinction is present in your experiment, then you generally want to refine the EXTI parameter. If this refines to 0, then, no, there is no point in it. But if it refines to a value around 3 times larger than the e.s.d, then you should reep it in the refinement. This will make a big difference for your R factorsm, but is especially noticeable in the residual density map.

Question 131. What do you mean working with merged datasets should not be done? Do you mean merging reflections by symmetry or datasets?

Answer. The data reduction should always return un-merge, un-cut reflection information.

Question 132. Mercury from CCDC would be one of the software you can use to visualize the symmetry elements?

Answer. Indeed.

Question 133. Right now it always show error warnings in Olex2 when you want to refine the structure as a twin using HKLF 5 file after you use Mask function. And some times the opposite wat also brings errors (refining the structure as a twin using HKLF 5 file and then use Mask function). Is Olex2 going to solve this problem in the future's upgrade?

Answer. I think you *are* using an old version. You can also always use PLATON (after LIST 8) for twinned structures.

Question 134. Is it possible to calculate the void using olex2?

Answer. Yes, there are different ways of calculating the voids -- it's under Tools > Maps

Question 135. How the max peak be taken care of?

In my structure I have seen only one counter anion. But it has to be two?  How can I proceed? Please suggest.

Answer. 'refine 4 30' will refine four cycles and show 30 peaks.

Question 136. How to get shelexT?

Answer. https://shelx.uni-goettingen.de/download.php

Question 137. (1) Missing FCF Refl between Thmin & STh/L= 0.600             6 Report

Answer. This means that there are reflections up to the resolution limit (Sin theta/lambda = 0.6) are missing. For example, it may be that you have a reflection under the beamstop that was not measurable. You are more likely to see this for lower symmetry but the answer to this is to extend the data collection strategy to measure the missing data.

Question 138. (1) Missing FCF Refl between Thmin & STh/L= 0.600             6 Report. How to fix this? (2) Large K Value in the Analysis of Variance ...... 4.131 Check.

Generally getting these two alerts in almost every structure?

Answer. This means there is an unusual statistical spread in your data. There are numerous reasons this could happen. Potentially the data have not been measured with long enough exposure times. It is a little difficult to diagnose without the data from your system.

Question 139. How to reduce Rint value?

Answer. It is a bit of an open ended question. First we need to know why Rint might be high. The most likely cause is that you have data that does not have a good enough I/sigma. The shorter your exposure times, the more likely the I/sigma is too low and the intensities are less reliable. Try running longer experiments, in particular with longer exposure times.

Question 140. Sometimes when I check the finalized cif file in checkCIF, it shows warnings "PLAT950_ALERT_5_G Calculated (ThMax) and CIF-Reported Hmax Differ", "PLAT955_ALERT_1_G Reported (CIF) and Actual (FCF) Lmax Differ by", and "PLAT992_ALERT_5_G Repd & Actual _reflns_number_gt Values Differ by". But during finalization process of cif file, I did not manually edit and change any related values of cif file. So can I know what is going on with my cif file? Why did I get these warnings?

Answer. The most common reason here is that there are rounding errors or rather differences between your refinement program and platon. This alert is marked as G which means you can check it but normally this is not an issue.

Question 141. How can we remove the residual density alert, already we assigned the solvent but still this residual density alert is coming. After squizing the solvent also it is coming?

Answer. Again this would depend on the reasons behind the residual density being high. The most common reasons are missed atoms, incorrect atoms, and fourier truncation peaks which are typically associated with heavily absorbing elements. Check your atoms are correct. In the last case they are usually symmetrically arranged on either side of the heavy element. Try checking your absorption correction. This error is far less likely when using the Analytical or Numerical methods for absorption correction as long as they are done properly.

Question 142. I have a structure which has large number of water molecules. Even after adding hydrogens one by one to these water molecules, I get alerts b, a) Short Inter D-H..H-D, b) D-H Bond without Acceptor. In this case, is it necessary to fix these hydrogens using dfix, or is there any other way to do it, I have already tried had-r and had commands?

Answer. You may need to build up your hydrogen bonding network more carefully -- it helps to check for symmetry equivalence atoms nearby.

Question 143. Lately I have had this alert, "PLAT142_ALERT_4_G s.u. on b - Axis Small or Missing," it is Alert G, but I don't know how to prevent it from appearing.

Answer. Is your space group correct? This could happen when an axis is equivalent by processing but later symmetry was changed to lower and there is no esd in the refinement files. I would expect this is unlikely to happen if only programs handle the files. Is the esd actually missing?

I also always see this alert. It seems the CrysAlisPro can yield unit cell parameters with very small esd. While as a user, I think this is a good thing, but it seems checkCIF does not "believe" the esd can be so small...

Actually this one is the other case where we have reflections beyond our resolution limit being missing. This is normal for area detector experiments since we cannot easily collect data only up to a given limit and not beyond it. Note that it is a G alert so you do not need to worry too much. Just check it is OK and acknowledge it. It would be upgraded to a C or higher alert if it is serious enough.

Question 144. How can we reduce the I/Sigma? sometimes alert comes corresponding to it? how we can clear it?

Answer. Reducing I/sigma is a bad idea. You want to increase it either by increasing exposure times or collecting more redundancy.

You need a longer exposure time. There is nothing you can do 'after the event' -- just re-collect new data, making sure to choose a suitably long exposure time. The pre-experiment will help with this.

Question 145. While merging data collected from the different crystals (crystals of the same compound obviously) does Rint usually increase?

Answer. Potentially. You may have crystals of different sizes which can have consequences for absorption correction, exposure times and so on. You therefore may end up with slightly different behaviour in each crystallite. There are no real yes no answers in this though.

Question 146. The 0.84 Angstrom resolution threshold is valid for both CuKa and MoKa sources or just for CuKa?

Answer. For copper, for Mo it is 0.75. For Mo 0.79 Angstrom can also be applied

The minimum required resolution is independent of the radiation, and it's always 0.84 ANG. For Mo, this is 50 degrees in 2 Theta, and in Cu it is ~135 degrees in 2 Theta.

Question 147. It seems checkCIF uses the rule of I/sigma > 2 rather than 3. So if high resolution srea data with I/sigma <3 were cut off by "OMIT", checkCIF will show a alert "PLAT909_ALERT_3_G Percentage of I>2sig(I) Data at Theta(Max) Still ", which warns some observed data, rather than the noise, was removed improperly. Should we take care of this warning?

Answer. G alerts are the least serious. If you know what you are doing it is up to you how to handle this. My general advice would be to not throw away data unless you absolutely have to. Simply lowering your R1 from e.g. 6% to 5% is not really a good reason to do this on its own. Protein crystallographers will use data down to I/sig =1 for example.

Question 148. PLAT230_ALERT_2_C Hirshfeld Test Diff for O00A --N00Z'' How can we remove this alerts?

Answer. Try to restraints the thermal parameter using SIMU or RIGU.

It means that two atoms in a bonding pair vibrate differently, which is not physically sensible.

Check the atom types are correct. Check occupancies are correct. Check ADP behaviour is not unexpected. Then use restraints if everything checks out.

Question 149. Apart from the known Acta Crystallographica Sections E and C, are there any other journals that you could suggest which focus mainly on synthesis and crystallography?

Answer. It depends on what you are doing. Maybe CrystEngComm.

There are quite a few journals. Here are a few examples:

https://www.rsc.org/journals-books-databases/about-journals/crystengcomm/

https://pubs.acs.org/journal/cgdefu

https://www.mdpi.com/journal/crystals

Question 150. what's the deal with "weighting scheme", I understand that you can check it in the olex2 program by clicking, but what is its essence?

Answer. https://www.iucr.org/resources/commissions/crystallographic-nomenclature/statdes/weight.html#:~:text=They%20estimate%20the%20variances%20of,refinement%20techniques%20are%20another%20example.

Question 151. PLAT430_ALERT_2_B Short Inter D...A Contact O007 ..O008Why we are getting this alerts and how can we remove?

Answer. Probably due to misplaced H atoms.

This is saying that you have a short hydrogen bond, check for disorder, and check hydrogen atoms are in the right place. You may need to manually place hydrogens to achieve proper placement

Question 152. How can we justify the high Rint (15-20%) in case of merged data? Is it enough to explain that it was not simply possible to collect full data from one crystal so the few data sets were merged? And that's what the high Rint is observed.

Answer. Generally everything can be justified to a referree, whether they will accept depends on the scientific basis for your justification. Please note that in CrysAlisPro we have tools to merge data from multiple crystals and the individual crystallite quality can be assessed with respect to the group. You may find some of your crystals have a poorer agreement or quality than the group. It may be advantageous to remove crystals from the group that are a poor fit.

Question 153. Is the high resolution structure gotten from high angle diffraction? and that is determined by the quality of the crystal?

Answer. High resolution means high theta. This is the axis around which the detector rotates. Higher resolution gives you smaller and smaller details in the data. E.g. Electron density requires very high resolution data. Your ability to measure high resolution may be affected by crystal quality or temperature. Low temperature data collection usually helps us achieve higher resolution

Question 154. i have collected an data using Mo source and it shows NPD. What should be done next? I have used low temperature also. yes,

Answer. Hard to answer without more details. I can only offer some pointers. Are all of the atoms in the structure correct? Have you missed a very heavy atom? Is the crystal diffracting to high resolution?

Question 155. all atoms in the structure are correct.

Answer. From my experience, for many coordinated complexes with a metal centre (the heaviest atom in the structure) I sometimes suffered NPD for this heaviest atom while using 100K but it behaves well while using 173K instead, just FYI.

May be the scaling problem either.

Question 156. What type of restrains and constrains are helpful to try and stop the hydrogens from rotating in your last example?

Answer. You could try a restraint which uses a known hydrogen bond acceptor, if you cannot find a way to restrain it, use AFIX 33 instead of AFIX 137. AFIX 33 will not allow rotation and place in idealised positions.

Question 157. For case study 5, why can't both Cl atoms are on the same C atom instead of 1 Cl and 1 F atoms?

Answer. The model is describing disorder. It is not that there is one molecule with both Cl and F bonded to that carbon. We have two molecules superimposed on each other. One molecule has F on the left, and Cl on the right and the other has Cl on the left and F on the right. I hope that makes sense. Imagine that alternating unit cells have one of those two molecules and the next would have the other.

Question 158. How voids and solvent accessible voids are different from each other?

Answer. It is called a solvent accessible void because it is large enough to accommodate a solvent molecule. There will be small voids in all structures but they are small and even H2O would not fit in there.

Question 159. So can we ignore voids?

Answer. If you get an alert from Platon, you need to explain it. So no.

Question 160. How the alert max peak 6.2 could be removed from check cif of an inorganic complex?

Answer. Yes -- ALWAYS do a proper face absorption correction!

Question 161. On one of the slides on resonant scattering, there is a flow chart at the bottom-right reading "anomalous dispersion, anomalous scattering, inelastic scattering, and resonant scattering". Are these all different phenomena? Do the arrows between the circles (going from left to right) have any meaning, in the sense of "leads to", "occurs prior to", "is larger than"?

Answer. No, they are just different names used by different fields of science and depending on the text-book. The last two can have more general meaning if you don't talk about the x-rays only

Question 162. What cause, or what should be done with, the alert: wR2 reported and calculated differ?

Answer. Can you provide the alert number for this? Without it is hard to guess what can be a problem.

Have you checked CIF with or without HKL data? In the latter case, make sure that your CIF matches the submitted FCF file or that the FCF block in CIF has been updated.

In general, that means the value you put in CIF does not match the one that check CIF can calculate from submitted data.

Question 163. From previous lecture of 3D-ED, data from a single grain will probably not guarantee high enough completeness. Is this going to negatively affect the determination of the absolute configuration?

Answer. Yes, but it is not as bad as you could think. I will show some case studies later in the lecture. You can always replace the completeness with resolution. Data is data; you just need enough of it.

You can also load several data sets into Jana2020 to increase completeness, you just cannot merge them beforehand. Jana can then refine the structure on all the imported datasets simultaneously.

Completeness is especially important in the initial solution step. Specifically for absolute structure, it is not that critical, as long as sufficiently many reflections recorded along zones with strong dynamical effects (i.e., those with many excited reflections) are included in the refinement.

Question 164. So, in the Experiment Strategy, is 'Chiral only' then indicating use of Sohncke space groups only?

Answer. Yes.  You can select from all non-centrosymetic groups and Sohncke SG. The former will also allow polar space groups.

Question 165. How can we change flack parameter?

Answer. It will be printed out at the last cycle of refinement if you use shell or olex2refine. Most of the programs will include it automatically in the summary of the refinement of non-centrosymmetric groups. In Olex2 you will notice an extra box in the section where you have Rint R1 I/sig etc.

Question 166. What is cu source of radiation hope it’s some wavelength of light on crystal. Please explain?

Answer. Sorry, I might use crystallographer slang. I meant an X-ray source with the copper target, which is a mixture of three wavelengths, actually A1    1.54056, A2, 1.54439  B1    1.39222, but the beta is removed mainly by mirrors in modern MF sources and very weak.

Question 167. Can the change in scale factor with frame number be an indication of beam damage/sample degradation?

Answer. Yes typically manifests as a consistent trend upwards in the scale factor as the diffracted intensity decreases

Question 168. A non merohedrical twin can only be generated by rotation?

Answer. Not always - sometimes the twin can be described as a mirror operation. Also, a non-centrosymmetric material can twin to look centrosymmetric - a special case of a merohedral twin.

Question 169. It is possible to differentiate racemic twins in the reciprocal space?

Answer. Existence of a racemic twin is not immediately obvious from the diffraction data, but rather is seen in the behaviour of the refinement (especially in the Flack or Parson's parameter). But there is a statistical test (cumulative intensities) that may suggest racemic twinning before structure solution.

Question 170. If I have two crystals with the same a and b cell parameters but with different c length and similar composition (probably trigonal or monoclinic platelets) is it possible to describe it as a twin or would it be something else?

Answer. This sounds like a "multi-crystal" if the two materials are chemically different. There could be other situations, even possibly incommensurate/modulated structures. This would require more detail about the material and the diffraction observations.

Question 171. I think the Rint was smaller in the integration process

Answer. Correct

Question 172. Is this issue of twining same as some miller indices value missing (hkl) that does not appear on the PXRD peak. For example a crystal plane of 110 which is h^2 + k^2 + l^2  = 6 but we cannot get the sum of square of three integers to be 7 or 15 because the crystal plane are siting above each other, thus the reason the peak on (2 theta degree) axis do not appear?

Answer. PXRD is essentially a twin of thousands/millions of components; all the intensities at a given resolution are collapsed onto 2theta. But I think you are referring to systematic absences.

Question 173. From hkl4 data, applying platon-twin in olex2, can one find a matrix for a non-merohedral twin? My interest in knowing if TWINROTMAT can be used to find a racemic twin?

Answer. Yes -- you can use either the Olex2 twin finding routine or TWIN ROTMAT to generate the hklf5 format file from within Olex2.

Question 174. Why didn't you try to solve the structure from the hklf5 file? Is it possible?

Answer. It is possible, but there is a lot of error in the overlapped reflections. Before solving the structure there is no model from which to calculate Fc, so we don't know the contribution each component makes to an overlapped reflection. Thus it's usually a better chance of solving the structure using only non-overlapped reflections.

Question 175. If the Rint is a lot higher in olex because of a problem with how the software interprets your data but it's lower in Crysalis, can you change the Rint so that it matches the value from Crysalis? Okay thanks! But then is the olex value also incorrect?

Answer. At the point of integration, an R int value is available, but that is 'lost' if you've only go the hklf5 format file.

We *could* simply print the value from the cif_od -- but is that really correct?

Question 176. Sometimes when I switch from HKLF 4 to HKLF 5 file, the Rint in Olex2 disappears (no value is shown). what should I do in this case? When I use  HKLF 5 file to rfine the structure and create cif file, the "_diffrn_reflns_number " and the "_reflns_number_total" become the same value. Why? How should I solve this problem?

Answer. R_int in twins is a really difficult concept. We are still not sure how to display these values. R_int of *which* component? We simply don't have the 'traditinal' R_int value from the intensity measurement alone.

Question 177. This example in olex the hkl was an hkl4? And you generate inside olex the hklf5?

Answer. Yes, initially it was a hkl4. After checking in OLEX2 if it was a twin (and apparently it was), OLEX2 found the twin matrix and generated a hkl5 file

Question 178. Following on for Horst: when I try this Twin tool in Olex2, I usually get the message ERR: Please chedk 'tools.get_twin_law_from_hklf5'. Is there an easy way around this?

Answer. This is strange -- and looks like a bug. Can you send me the 'log' file after the error happens?

Question 179.

Answer.